Risk Management

Revision History

| Rev | Date | Revision History |
| --- | --- | --- |
| 00 | 2025.02.28 | First Issued. |

Contents

1.	Scope and Purpose	3

1.1	Purpose	3

1.2	Scope	3

1.3	Software Identification	3

1.4	Reference	3

2.	Risk Management Plan	3

3.	Risk Analysis	4

3.1	Risk analysis process	4

3.2	Intended use and reasonably foreseeable misuse	6

3.3	Identification of characteristics related to the safety of the medical device	6

3.4	Identification of hazards and hazardous situations	17

3.5	Risk Estimation	17

4.	Risk Evaluation	18

5.	Risk Control	18

5.1	Risk control option analysis	18

5.2	Implementation of risk control measures	18

5.3	Residual risk evaluation	18

5.4	Benefic-risk Analysis	19

5.5	Risk arising from risk control measures	19

5.6	Completeness of risk control	19

6.	Evaluation overall residual risk	19

7.	Risk management review	19

- Scope and Purpose

  - Purpose

This document provides a description of software risk management.

  - Scope

This document applies only to software hazards in the MCU. It does not include hazards to the hardware.

  - Software Identification

| Software name | FasterDx |
| --- | --- |
| Software version | 1.00.00 |
| Operating system | Windows 10 IoT Ent. |
| Programmable language | C++, C# |
| Safety Level | Class B |

  - Reference

| Reference | Title of standard |
| --- | --- |
| 1) ISO 14971 [2019] | Medical devices –Application of risk management to medical devices. |
| 2) IEC 60601-1 [2005]+AM1[2012] | Medical electrical equipment – General requirements for basic safety and essential performance |
| 3) IEC 60601-1-2 [2014] | Medical electrical equipment. General requirements for basic safety and essential performance. Collateral Standard. Electromagnetic disturbances. Requirements and tests. |
| 4) ISO 7010:2011 | Graphical symbols -- Safety colors and safety signs -- Registered safety signs |
| 5) IEC 60601-1-6 [2010]+AMD1[2013] | Medical Electrical Equipment - Part 1-6: General Requirements for Basic Safety and Essential Performance - Collateral Standard: Usability |
| 6) IEC 62366 Ed. 1.0 [2007] +AMD1[2014] | Medical devices - Application of usability engineering to medical devices |

- Risk Management Plan

- Software safety levels are documented in the Software Requirements

- The theoretical basis for the use of low-level classification and the results of each documented risk management method implementation verification and verification are documented in the software design specification.

- Implementation of each documented risk management method must be verified, and the verification results are documented in the software Verification, Validation and Test Report.

- RISK MANAGEMENT includes PEMS, and therefore IEC 60601-1, Clause 14 must be applied as a requirement. Clause 14 applies to this device because it is difficult to demonstrate through ISO 14971 activities that failure does not pose an unacceptable risk.

- ISO 14971 risk management activities are for the entire product life cycle, from product planning to development, production, and sales.60601-1 Clause 14 is a requirement for PEMS and IEC 62304 was applied as the software development process of this device to effectively satisfy this requirement. IEC 62304 was applied to manage hazards that may arise from omission of development procedures.

- Architecture design is performed according to the Software Requirements Specification (SRS) of this product, safety classification is defined, and possible risks are identified. The development process according to IEC 62304 is applied according to the defined classification.

- The Risk control method is performed in one or more of the following methods, and the risk control method for the identified risk is reflected in the architectural design and software design specification (SDS)

- Integrity Parts

- Failure - Safety Features

- Redundancy

- Diversity

- division of functionality

- Risk Acceptance criteria: The risk acceptance criteria for the first identified risk and residual risk after risk control follow the acceptability of risk

- Risk Analysis

  - Risk analysis process

Risk assessment method for each risk factor: Analyze considering probability and severity.

| Level | Risk severity | Severity | Example |
| --- | --- | --- | --- |
| 5 | Catastrophic / Fatal | Results in death | Death |
| 4 | Critical | Results in permanent impairment or irreversible injury | Patient tissue damage, bleeding |
| 3 | Serious / Major | Results in injury or impairment requiring medical or surgical intervention | Infection, electric shock, burns |
| 2 | Minor | Results in temporary injury or impairment not requiring medical or surgical intervention | Delayed treatment |
| 1 | Negligible | Results in inconvenience or temporary discomfort | Discomfort during use (product damage, product quality deterioration) |
| * Reference of definition; ISO/TR 24971:2020 Clause 5. |  |  |  |

| Level |  | Possibility of occurrence |
| --- | --- | --- |
| 5 | Frequent | ≥ 10-3 |
| 4 | Probable | < 10-3 and ≥ 10-4 |
| 3 | Occasional | < 10-4 and ≥ 10-5 |
| 2 | Remote | < 10-5 and ≥ 10-6 |
| 1 | Improbable | < 10-6 |
| * Reference of definition; ISO/TR 24971:2020 Clause 5. |  |  |

- 2 Hazardous Areas

It uses a "metric method" that combines probability and severity. All calculated risks are considered by dividing them into three risk areas. Each risk area is controlled differently according to the level of risk based on ISO/TR 24971:2020.

- How to calculate the risk

Risk Level = Possibility x Severity

- Risk Matrix

| Frequent (5) | 5 | 10 | 15 | 20 | 25 |
| --- | --- | --- | --- | --- | --- |
| Probable (4) | 4 | 8 | 12 | 16 | 20 |
| Occasional (3) | 3 | 6 | 9 | 12 | 15 |
| Remote (2) | 2 | 4 | 6 | 8 | 10 |
| Improbable (1) | 1 | 2 | 3 | 4 | 5 |
| Possibility Severity | (1) Negligible | (2) Minor | (3) Serious | (4) Critical | (5) Catastrophic |

- Green Zone (1 ~ 4) : ACC (Acceptable)

- Yellow Zone (5 ~ 9) : AFAP (As Far As possible)

- Red Zone (10 ~ 25) : N/ACC (Non-Acceptable)

  - Intended use and reasonably foreseeable misuse

Raywatt Inc.'s high-speed cardiovascular OCT system utilizes dedicated software (FasterDx) to perform optical coherence tomography (OCT) of coronary arteries. This software generates high-resolution images of coronary arteries using Raywatt Inc.'s OCT system and compatible imaging catheters, allowing for the quantitative analysis of lesion characteristics.

Medical professionals can use this software to precisely evaluate the structure of the vascular endothelium, plaque composition, and the degree of vascular stenosis, which aids in developing treatment plans. The high-resolution OCT imaging assists in evaluating vascular conditions before and after stent placement and serves as a guide during procedures.

This product is not considered for misuse because it is used by doctors or related experts.

  - Identification of characteristics related to the safety of the medical device

- Identification of hazards and characteristics related to safety (ISO/TR 24971:2020 Annex.A2 Questions)

- The Analysis of the safety factors are described in the Device Risk Management Report

| No. | ISO/TR 24971:2020 Annex.A2 Questions | A / N | Identified Characteristics |
| --- | --- | --- | --- |
| A.2.1 | What is the intended use and how is the medical device to be used? | A | Device for measuring the internal structure of the cardiovascular disease areas using lasers. |
| A.2.2 | Is the medical device intended to be implanted? | N/A | This product is not a life-sustaining device. |
| A.2.3 | Is the medical device intended to be in contact with the patient or other persons? | A | The product itself is free of patient contact. When using the Chouette OCT catheter, Contact time: ≤ 3 minutes (per use) Frequency of contact: Continuous contact during use (insertion into the circulatory blood vessel) |
| A.2.4 | What materials or components are utilized in the medical device or are used with, or are in contact with, the medical device? | A | Rotary Junction cover (Surgical drapes to prevent contamination caused by placing the rotating joints on the patient's C-arm bed during product use.) Chouette OCT catheter |
| A.2.5 | Is energy delivered to or extracted from the patient? | A | Optical energy |
| A.2.6 | Are substances delivered to or extracted from the patient? | N | This product does not supply or extract substances |
| A.2.7 | Are biological materials processed by the medical device for subsequent reuse, transfusion or transplantation? | N | This product does not handle biological substances |
| A.2.8 | Is the medical device supplied sterile or intended to be sterilized by the user, or are other microbiological controls applicable? | N | This product is not provided sterilized |
| A.2.9 | Is the medical device intended to be routinely cleaned and disinfected by the user? | N | This product does not require routinely cleaning |
| A.2.10 | Does the medical device modify the patient environment? | N | This product does not alter the patient's environment |
| A.2.11 | Are measurements taken? | A | Measured variables: include the diameter of blood vessels, characteristics of plaque, stent placement, and tissue properties. |
| A.2.12 | Is the medical device interpretative? | N | Not applicable. |
| A.2.13 | Is the medical device intended for use in conjunction with other medical devices, medicines or other medical technologies? | A | Receive Angio images if necessary |
| A.2.14 | Are there unwanted outputs of energy or substances? | A | Noise, vibration, electromagnetic waves |
| A.2.15 | Is the medical device susceptible to environmental influences? | N | This product is not sensitive to environmental influences. |
| A.2.16 | Does the medical device influence the environment? | N | Not applicable. |
| A.2.17 | Does the medical device require consumables or accessories? | A | Chouette OCT catheter and Rotary junction cover |
| A.2.18 | Is maintenance or calibration necessary? | A | Maintenance or calibration is necessary |
| A.2.19 | Does the medical device contain software? | A | This product is a medical device, including software, and software installation/validation/modification must be performed only by experts. |
| A.2.20 | Does the medical device allow access to information? | A | Export of OCT images measured via Ethernet port and USB ports is possible. |
| A.2.21 | Does the medical device store data critical to patient care? | A | Patient data is encrypted and stored. |
| A.2.22 | Does the medical device have a restricted shelf life? | A | In the case of catheters, there is an expiration date. |
| A.2.23 | Are there any delayed or long-term use effects? | A | This product is not a medical device with delayed/long-term effects. |
| A.2.24 | To what mechanical forces will the medical device be subjected? | N | This product is a medical device that does not include mechanical forces. |
| A.2.25 | What determines the lifetime of the medical device? | A | Expected service life of the light source is 10 years |
| A.2.26 | Is the medical device intended for single use? | N | This product is a reusable device. |
| A.2.27 | Is safe decommissioning or disposal of the medical device necessary? | N | Not applicable. |
| A.2.28 | Does installation or use of the medical device require special training or special skills? | A | This product is a medical device, including software, and software installation/validation/modification must be performed only by experts. |
| A.2.29 | How will information for safety be provided? | A | Safety information is provided through the user manual. |
| A.2.30 | Are new manufacturing processes established or introduced? | N | Not applicable. |
| A.2.31 | Is successful application of the medical device dependent on the usability of the user interface? |  |  |
| A.2.31.1 | Can the user interface design features contribute to use error? | N | Not applicable. |
| A.2.31.2 | Is the medical device used in an environment where distractions can cause use error? | N | Not applicable. |
| A.2.31.3 | Does the medical device have connecting parts or accessories? | A | Chouette OCT catheter and Rotary junction |
| A.2.31.4 | Does the medical device have a control interface? | A | This product provides handles, wheels, and software (FasterDx) GUI for users to control and move the device. |
| A.2.31.5 | Does the medical device display information? | A | This product displays information such as video taken through software and setting option. |
| A.2.31.6 | Is the medical device controlled by a menu? | A | This product is controlled by menu which provides to enable photographing and control through a device connected through an interface of the software |
| A.2.31.7 | Is the successful use of the medical device dependent on a user’s knowledge, skills and abilities? | A | It can only be used by experts with relevant training. |
| A.2.31.8 | Will the medical device be used by persons with specific needs? | A | It can only be used by experts with relevant training. |
| A.2.31.9 | Can the user interface be used to initiate unauthorised actions? | N | Unauthorized users are not allowed to make any modifications to the product device without the manufacturer's permission. |
| A.2.32 | Does the medical device include an alarm system? | A | Alarm window pop-up in case of malfunction. |
| A.2.33 | In what ways might the medical device be misused (deliberately or not)? | N | Not applicable. |
| A.2.34 | Is the medical device intended to be mobile or portable? | A | It is a portable device. |
| A.2.35 | Does the use of the medical device depend on essential performance? | A | It depends on performance such as light source and motor. |
| A.2.36 | Does the medical device have a degree of autonomy? | N | Users can select and perform appropriate corrective actions. |
| A.2.37 | Does the medical device produce an output that is used as an input in determining clinical action? | N | Not applicable. |

- Identification of known and foreseeable HAZARDS in PEMS (IEC60601-1 14.6.1)

| IEC 60601-1 14.6.1 Lists | A / N | Identified Characteristics |
| --- | --- | --- |
| undesired feedback [physical and data] (possibilities include: unsolicited input, out of range or inconsistent input, and input originating from electromagnetic interference); | A | In the present system, an image of an unintended section may be input by physical manipulation when photographing an OCT image. |
| unavailable data | N | Not applicable. |
| lack of integrity of data | N | Not applicable. |
| incorrect data | N | Not applicable. |
| incorrect timing of data. | N | Not applicable. |
| unintended interactions within and among PESS | N | Not applicable. |
| unknown aspects or quality of third-party software | N | Not applicable. |
| unknown aspects or quality of third-party PESS; | N | Not applicable. |
| lack of data security, including its effects on data privacy, and particularly vulnerability to tampering, unintended interaction with other programs and viruses. | N | Not applicable. |
| failure of the IT-NETWORK to provide the characteristics necessary for the PEMS to achieve its BASIC SAFETY or ESSENTIAL PERFORMANCE. | N | Not applicable. |

- Causes of HAZARDOUS SITUATIONS associated with IT-NETWORKS (IEC60601-1 H.7.2)

| IEC 60601-1 H.7.2 Lists | A / N | Identified Characteristics |
| --- | --- | --- |
| remote servicing (external access to the network); | N | Not applicable. |
| operating system (compatibility of operating systems) | N | Not applicable. |
| modification/upgrades of software (operating systems, applications, etc.) | A | Software upgrades are carried out by the manufacturer if necessary. |
| interface compatibility (data collisions, data formats) | N | Not applicable. |
| connections (modification of hardware, network connectors) | N | Not applicable. |
| network interface boards (compatibility) | N | Not applicable. |
| network protocols (DICOM, HL7, etc.) | N | Not applicable. |
| packet address structure/timing | N | Not applicable. |
| normal network loads/bandwidth | N | Not applicable. |
| peak network load | N | Not applicable. |
| data media (longevity and retrievability) | N | Not applicable. |
| security (viruses, worms, unauthorized software updates or upgrades) | N | Not applicable. |
| maximum acceptable response time | N | Not applicable. |
| acceptable failure rate of the network | N | Not applicable. |
| availability of the network (planned and unplanned maintenance) | N | Not applicable. |
| inconsistency in interfaces/formats resulting in loss of fidelity during information transfer | N | Not applicable. |
| heterogeneous network topologies | N | Not applicable. |

  - Identification of hazards and hazardous situations

The hazards identified in the "Hazard Analysis" are reflected in the requirements document, as shown in the table. In addition, after the requirements of PEMS and PESS were confirmed, the hazard situation was analyzed for each requirement. The results are listed in the " Refer to Attachment 1. FMEA Table".

  - Risk Estimation

A risk estimation of the hazard situation analyzed was carried out for each requirement, and the results are listed in the " Refer to Attachment 1. FMEA Table". The table below summarizes the results.

| Frequent (5) | 9 | 4 | 0 | 0 | 0 |
| --- | --- | --- | --- | --- | --- |
| Probable (4) | 0 | 5 | 0 | 0 | 0 |
| Occasional (3) | 0 | 0 | 0 | 0 | 0 |
| Remote (2) | 0 | 0 | 0 | 0 | 0 |
| Improbable (1) | 0 | 0 | 0 | 0 | 0 |
| Possibility Severity | (1) Negligible | (2) Minor | (3) Serious | (4) Critical | (5) Catastrophic |

As a result of the initial calculation of each risk factor, risk reduction measures are carried out for all existing risks and residual risks are evaluated.

- Risk Evaluation

For each risk situation, a residual risk assessment was carried out after risk control, and the results are listed in the "Risk Management Report". The table below summarizes the results.

| Frequent (5) | 9 | 4 | 0 | 0 | 0 |
| --- | --- | --- | --- | --- | --- |
| Probable (4) | 0 | 5 | 0 | 0 | 0 |
| Occasional (3) | 0 | 0 | 0 | 0 | 0 |
| Remote (2) | 0 | 0 | 0 | 0 | 0 |
| Improbable (1) | 0 | 0 | 0 | 0 | 0 |
| Possibility Severity | (1) Negligible | (2) Minor | (3) Serious | (4) Critical | (5) Catastrophic |

The initial calculation value of each risk factor decreased after the implementation of risk control. In addition, since all of these levels have been directed to acceptable zones, an assessment of the total residual risk tolerance is deemed appropriate.

- Risk Control

In the "Risk assessment before risk control", the risk control for the hazardous situation identified for each requirement was carried out, and the results were listed in the "Attachment 1. FMEA Table".

  - Risk control option analysis

Refer to Attachment 1. FMEA Table

  - Implementation of risk control measures

Refer to Attachment 1. FMEA Table

  - Residual risk evaluation

For each risk situation, a residual risk assessment was carried out after risk control, and the results are listed in " Attachment 1. FMEA Table". The table below summarizes the results.

| Frequent (5) | 0 | 0 | 0 | 0 | 0 |
| --- | --- | --- | --- | --- | --- |
| Probable (4) | 0 | 0 | 0 | 0 | 0 |
| Occasional (3) | 0 | 0 | 0 | 0 | 0 |
| Remote (2) | 1 | 2 | 0 | 0 | 0 |
| Improbable (1) | 8 | 7 | 0 | 0 | 0 |
| Possibility Severity | (1) Negligible | (2) Minor | (3) Serious | (4) Critical | (5) Catastrophic |

The initial calculation value of each risk factor decreased after the implementation of risk control. In addition, since all of these levels have been directed to acceptable zones, an assessment of the total residual risk tolerance is deemed appropriate.

  - Benefit-risk Analysis

None

  - Risk arising from risk control measures

None

  - Completeness of risk control

All risk factors were identified, and a risk management report was prepared based on the results of the determination of the likelihood and severity of occurrence. As a result, all non-acceptance risks have become acceptable risks, and those that can be reasonably performed will be managed by the quality plan. In addition, all risk factors identified through risk analysis will be continuously monitored.

| Item | Total | ACC (Acceptable) | AFAP (As Far As possible) | N/ACC (Non-Acceptable) |
| --- | --- | --- | --- | --- |
| Risk Assessment Statistics | 18 | 0 | 14 | 4 |
| Statistics after risk control | 18 | 18 | 0 | 0 |

| Risks that require control measures after risk assessment | 0 |
| --- | --- |
| Risks that require continuous management after risk control | 0 |

- Evaluation overall residual risk

No residual risk

All residual risks are derived into the Green Zone (Broadly Acceptable Zone), so risk/benefit analysis is not required.

- Risk management review

The results of the risk management process have been reviewed to ensure that the risk management plan has been appropriately executed, that the overall residual risk is acceptable, and that appropriate methods are in place to collect and review relevant production and post-production information.

The risk management review was performed after implementation and verification of all risk control measures but prior to commercial release of the medical device.