Documentation Level Evaluation

- Device Name

- FASTER OCT Imaging System

- Description

- Description: The FASTER OCT Imaging System, used with the compatible Chouette Imaging Catheter, is intended for imaging of coronary arteries and is suitable for patients who are candidates for percutaneous interventional procedures. The software is designed for medical professionals involved in the diagnosis and treatment of cardiovascular diseases.

3. Rationale

- Rationale: A failure or latent flaw of the device software function(s) would not present a hazardous situation with a probable risk of death or serious injury to either a patient, user of the device, or others in the environment of use, prior to the implementation of risk control measures. The device is intended for use as an aid in the assessment of mild (non-severe) injury and is not intended as a standalone diagnostic.

4. Outcome

- Basic Documentation Level