Software Description

Revision History

| Rev | Date | Revision History |
| --- | --- | --- |
| 00 | 2025.02.28 | First Issued. |

Contents

1.	Software Operation	4

1.1	The role of the software as it relates to the intended use of the device	4

1.2	The intended user of the software	4

1.3	The intended patient population	4

1.4	The analysis methodology	4

1.5	The device software function uses ML models trained through ML methods	5

2.	Software Specifications	5

2.1	Hardware Platforms	5

2.2	Software Platforms	6

2.3	OTS (Off-the-Shelf) Software	6

2.4	Final Release Version	6

3.	Software Inputs and Outputs	6

3.1	The Inputs and Their Format	6

3.2	Who or What Provides the Inputs?	6

3.3	The Outputs and Their Format	7

3.4	Who or What Receives the Outputs?	7

3.5	Does This Software Replace or Affect Tasks Previously Performed Manually by Medical Professionals, Patients, Caregivers, or Other Operators?	7

3.6	Clinical Workflow (Step-by-Step Process and Preconditions)	7

3.7	Is the Device Designed to Be Interoperable?	8

- Software Operation

  - The role of the software as it relates to the intended use of the device

Raywatt Inc.'s high-speed cardiovascular OCT system utilizes dedicated software (FasterDx) to perform optical coherence tomography (OCT) of coronary arteries. This software generates high-resolution images of coronary arteries using Raywatt Inc.'s OCT system and compatible imaging catheters, allowing for the quantitative analysis of lesion characteristics.

Medical professionals can use this software to precisely evaluate the structure of the vascular endothelium, plaque composition, and the degree of vascular stenosis, which aids in developing treatment plans. The high-resolution OCT imaging assists in evaluating vascular conditions before and after stent placement and serves as a guide during procedures.

  - The intended user of the software

This software is designed for medical professionals involved in the diagnosis and treatment of cardiovascular diseases. The primary users include:

- Interventional cardiologists

- Physicians performing cardiovascular interventional procedures

- Medical personnel involved in cardiac catheterization and intravascular imaging

- Clinical researchers and scientists studying cardiovascular diseases

The software is used in professional medical environments (e.g., interventional cardiology labs, cardiac catheterization labs) and helps clinicians analyze patient lesions quantitatively for treatment planning.

  - The intended patient population

This software is intended for patients suspected of having or already diagnosed with coronary artery disease (CAD).

- Does the software focus on a specific disease, condition, patient characteristic, or demographic?
Yes, the software is specifically designed for patients with coronary artery disease. It is used to assess vascular stenosis associated with atherosclerosis, angina, and myocardial infarction.

- Does the software provide information directly applicable to a specific disease or condition?
Yes, the software provides quantitative analysis of coronary artery stenosis and lesion morphology. Based on OCT imaging, it measures plaque composition, lesion length, and vascular diameter, enabling medical professionals to assess lesion characteristics and determine the appropriate treatment approach (e.g., medication or stent implantation).

  - The analysis methodology

Raywatt Inc.'s high-speed cardiovascular OCT system software provides high-resolution visualization and analysis of coronary artery structures using optical coherence tomography (OCT). The software applies image data analysis techniques for quantitative lesion assessment.

OCT Image Analysis Techniques

- Image Preprocessing & Enhancement

  - Removes noise and enhances contrast for clearer OCT images

- Measurement of Vascular Diameter and Lesion Length

  - Quantitative analysis of vascular diameter and lesion length, aiding in interventional treatment decisions

AI and Machine Learning-Based Analysis

- Deep learning-based vascular segmentation and lesion auto-analysis

  - Uses Convolutional Neural Networks (CNNs) to automatically detect vascular contours, stents, vessel landmarks, and guidewires from OCT images

  - Detects calcium deposits for more precise lesion analysis

  - Applies real-time object detection and classification to rapidly and accurately identify key structures in OCT images

This software combines rule-based calculations, artificial intelligence (AI), and machine learning (ML) algorithms to enhance the accuracy and reliability of OCT image analysis, supporting rapid and precise decision-making by medical professionals.

  - The device software function uses ML models trained through ML methods

- Methods, models, frameworks, and/or platforms used

Uses deep learning-based image analysis models (CNN-based) to automatically detect and quantify vascular contours, stents, vessel landmarks, guidewires, and calcium deposits from OCT images.

- Data (populations, samples) used for model training

The model is trained using OCT image datasets collected from partner research institutions.

- How, when, and where was the data collected?

Time: Continually updated with the latest OCT images since the initial model development

Location: Data is collected from hospitals and research institutions through collaborative research

Method: OCT imaging data is processed and anonymized before being used for machine learning training

- Steps taken to identify and address potential biases and limitations of the models

A feedback loop is established to continuously monitor and improve the bias and limitations of AI models.

- Materials, mechanisms, and/or approaches used to ensure transparency in model development, performance, and limitations

Performance metrics (e.g., accuracy, sensitivity, specificity) are published and regularly validated.

- Software Specifications

  - Hardware Platforms

Main Board: [DFI] RPS630-R680E

CPU: [Intel] i7-12700

GPU: [GIGABYTE] GeForce RTX 4060 Ti WINDFORCE OC D6 8GB

Memory: [Samsung] DDR5 32GB PC5-44800 (5600) *2EA

Storage: [Samsung] 980 PRO M.2 NVMe 2TB

  - Software Platforms

Windows 10 IoT Enterprise

  - OTS (Off-the-Shelf) Software

PostgreSQL: Relational database for storing, searching, modifying, and deleting data

OpenCvSharp: Computer vision and image processing in the .NET environment

Json.NET: JSON data serialization and deserialization

log4net: Application execution logging and error tracking

Math.NET Numerics: Mathematical operations, matrix calculations, and statistical analysis

SharpVectors.Wpf: Renders and converts SVG (vector graphics) files in WPF environments

OpenCV: Image filtering, contour detection, and object tracking

VTK (Visualization Toolkit): 3D data rendering and visualization

DCMTK (DICOM Toolkit): DICOM medical image conversion, storage, retrieval, and PACS integration

Numpy: Multi-dimensional array operations and numerical calculations

Pillow: Image transformation, filtering, and format conversion

Scipy: Scientific computing, signal processing, optimization, and interpolation

Pywin32: Windows API access and system control

NVIDIA CUDA Toolkit: GPU acceleration for AI and image processing

  - Final Release Version

1.00.00

- Software Inputs and Outputs

  - The Inputs and Their Format

Input Data:

- OCT (Optical Coherence Tomography) image data

- Patient's basic data (name, date of birth, gender)

- Analysis parameters set by medical professionals (procedure, vessel type, vascular region)

Input Data Format:

- OCT Images: RAW data

- Patient Data: GUI input

- Analysis Parameters: GUI input

  - Who or What Provides the Inputs?

Medical Devices:

- Real-time image data is collected from OCT imaging equipment.

Medical Professionals:

- Manually input patient data

- Select the lesion to be analyzed and enter necessary settings.

  - The Outputs and Their Format

Output Data:

- Analyzed OCT images and 3D vascular structure information

- Quantitative analysis results of lesions (vascular diameter, area, lesion length, etc.)

- AI analysis results (vascular contours, stents, vessel landmarks, calcium deposits, etc.)

Output Data Format:

- Image Data: RAW, DICOM, PNG/JPEG/TIFF/MP4

- Numerical Data:

  - If the image is RAW: Stored in JSON format

  - If the image is DICOM: Stored within the DICOM file

  - Other formats: Embedded within the image or video

  - Who or What Receives the Outputs?

Medical Professionals:

- Interventional cardiologists, radiologists, and interventional procedure specialists

Medical Information Systems:

- Hospital PACS (Picture Archiving and Communication System) stores the imaging results.

  - Does This Software Replace or Affect Tasks Previously Performed Manually by Medical Professionals, Patients, Caregivers, or Other Operators?

Previously, medical professionals manually identified vascular contours and measured lesions in OCT images.

This software automates vascular contour detection, lesion length measurement, and calcium deposit identification using AI-based analysis.

Reduces workload for medical professionals and improves the objectivity and accuracy of analysis.

  - Clinical Workflow (Step-by-Step Process and Preconditions)

- Preparation for OCT Imaging

  - Obtain patient consent and insert the OCT catheter.

  - Inject contrast agent (or saline) and capture vascular images.

- OCT Image Acquisition

  - Obtain OCT images and process them.

- AI-Based Analysis and Result Generation

  - Analyze vascular contours, lesion location, and stent condition.

  - Perform AI-based lesion quantification and data visualization.

- Medical Review of Analysis Results

  - Medical professionals evaluate lesions based on automated analysis results.

  - If necessary, perform additional manual analysis.

- Diagnosis and Treatment Planning

  - Decide on intervention based on analysis results (e.g., whether to insert a stent).

  - Store final results in DICOM format within PACS.

- Preconditions:

  - OCT equipment and catheter must be prepared.

  - Is the Device Designed to Be Interoperable?

No network or wireless communication functionality is included.

The device does not transmit, exchange, or use information with other medical/non-medical devices, systems, or products via electronic interfaces.