Software Life Cycle Process Document

Revision History

| Rev | Date | Revision History |
| --- | --- | --- |
| 00 | 2025.02.28 | First Issued. |

Contents

1.	Introduction	4

1.1	Purpose	4

1.2	Scope	4

1.3	Reference	4

1.4	Definitions	4

2.	Software Life Cycle Overview	5

2.1	Software Safety Classification	5

2.2	System Requirements	5

2.3	Design Phases	6

2.4	Avoiding Common Software Defects Based on Selected Programming Technology	6

3.	Software Development Process	6

3.1	Software Development Planning	6

3.2	Software Requirements Analysis	7

3.3	Software Architecture Design	7

3.4	Software Detailed Design	7

3.5	Software Implementation	7

3.6	Software Testing	7

4.	Software Maintenance Process	7

4.1	Creation of Change Request	7

4.2	Evaluation of Change	8

5.	Software Risk Management Process	8

6.	Software Configuration Management Process	9

7.	Software Problem Resolution Process	9

8.	Documentation and Record Management	11

8.1	Software Life Cycle	11

8.2	Activities at Each Software Life Cycle	11

Attachment - Software – IEC 62304:2006+AMD1:2015	13

- Introduction

  - Purpose

This document defines the software development process and activities. To ensure software quality, this procedure describes the requirements for software design control, verification, and validation.

  - Scope

This document aims to integrate and support general design control requirements. Project managers, quality assurance managers, and product managers shall use this plan as a tool to manage software development activities throughout the project lifecycle.

  - Reference

- IEC 62304:2006/A1:2016 Medical device software — Software life cycle processes

- ISO 13485:2016 Medical devices — Quality management systems — Requirements for regulatory purposes

- ISO 14971:2019 Medical devices –Application of risk management to medical devices.

  - Definitions

- Software System

A set of software items, associated documentation, and data that form an integrated whole.

- Software Item

A part of the software system that is separately defined and managed.

- Software Unit

The smallest independently testable part of a software item.

- Software Development Life Cycle

The overall process of software development, including requirements analysis, design, implementation, verification, and maintenance.

- Software Safety Classification

The categorization of software is based on the potential harm it may cause to patients, users, or the environment in the event of failure:

- Class A: No injury or damage to health is possible.

- Class B: A non-serious injury is possible.

- Class C: Death or serious injury is possible.

- Software Verification

The process of evaluating whether the software meets specified requirements.

- Software Validation

The process of ensuring that the software performs as intended for its specified purpose.

- Software Maintenance

Activities performed after the software has been released to improve performance, fix defects, or adapt to changes in the environment.

- Software Configuration Item

A specific version or component of the software that is subject to configuration management.

- Software Risk Management

The process of identifying, assessing, and controlling risks associated with software.

- Software Life Cycle Overview

  - Software Safety Classification : Class B

Programming Languages

| Name | Version |
| --- | --- |
| C++ | 14 |
| C# | 10.0 |

Development Software

| Name | Version |
| --- | --- |
| Microsoft Visual Studio | 2022 |

  - System Requirements

Target Runtime:

| Name | Version |
| --- | --- |
| .NET | 6.0 |

System specification:

- Main Board: [DFI] RPS630-R680E

- CPU: [Intel] i7-12700

- GPU: [GIGABYTE] GeForce RTX 4060 Ti WINDFORCE OC D6 8GB

- Memory: [Samsung] DDR5 32GB PC5-44800 (5600) *2EA

- Storage: [Samsung] 980 PRO M.2 NVMe 2TB

  - Design Phases

| Activity | Responsibilities | Review method |
| --- | --- | --- |
| Planning | Software engineer | RW-TF-FST-02-04(00) Software Requirements Specifications (SRS) |
| Implementation | Software engineer | RW-TF-FST-02-05(00) Software Architecture Description RW-TF-FST-02-06(00) Software Design Specifications (SDS) |
| Test | Software engineer, Test engineer | RW-TF-FST-02-08(00) Software V&V Report |
| Risk management | Quality manager, Test engineer | RW-TF-FST-02-03(00) Risk Management |
| Release | Software engineer | RW-TF-FST-02-09(00) Software Version Revision Level History |

  - Avoiding Common Software Defects Based on Selected Programming Technology

Test and peer code review can help identify and correct potential defects before they become problematic.

Defects based on Selected Programming Technology:

| Programming Technologies | Potential defects | Prevention strategies |
| --- | --- | --- |
| C++ | Dynamic memory allocation may lead to access violation errors or memory leak in runtime. | Use static analysis tools for testing and use smart pointers. |
| C++ | Buffer overflow | Use standard libraries like std::vector or std::string. |
| C++, C# | Null reference exceptions | Always initialize pointers or object before use and perform null pointer checks. |

- Software Development Process

  - Software Development Planning

- Input

- Software Requirements Analysis

- Software Risk Analysis

- Output

- Software Maintenance Process

- Software Risk Management Process

- Software Configuration Management Process

- Software Problem Resolution Process

- Documentation and Record Management

  - Software Requirements Analysis

- Input

- System Requirements Analysis

- Software Risk Analysis

- Output

- RW-TF-FST-02-04(00) Software Requirements Specifications

  - Software Architecture Design

- Input

- System and Software Requirements Analysis

- Risk Analysis Risk Analysis

- System Structural Design

- Output

- RW-TF-FST-02-05(00) Software Architecture Description

  - Software Detailed Design

- Input

- System Requirements Analysis

- Software Risk Analysis

- Output

- RW-TF-FST-02-04(00) Software Requirements Specifications

- RW-TF-FST-02-08(00) Software V&V Report

  - Software Implementation

For each pull request, a code review is performed by a team member who was not the main author of the code under review. The code review is only marked as “approved” if the code complies with the code review criteria. This is:

- Code fulfils the software requirements.

- Code is built without error.

- No conflicts with other code.

- Adherence to C++ at google style guide, C# at google style guide.

  - Software Testing

Refer to RW-TF-FST-02-08(00) Software V&V Report

- Software Maintenance Process

  - Creation of Change Request

Changing proposals can originate from various sources, e.g.:

- Internal design ideas

- Market research

- Customer feedback (see process for feedback management)

- Post-Market Surveillance (see process for post-market surveillance)

- Changes to the QMS

They can originate from anywhere inside the company.

The Product Manager creates a change request activity in Redmine with a description of the proposed change.

Participants

Anyone in the company with a change proposal

| Input | Output |
| --- | --- |
| Change proposal | Documented change request (ID per product version) |

  - Evaluation of Change

The Product Manager, together with a Regulatory Affairs Manager and the Head of Software Development, then evaluates the proposed change. In the first step, it is assessed if the change constitutes a bug fix or not. In case of a bug fix, the Product Manager directly proceeds with the evaluation as part of the bug fixes documentation list.

If the change does not constitute a bug fix but rather a regular change request, it is evaluated whether the change:

- Is a significant change

- Introduces any new risks and whether those are acceptable

- Impacts any existing risk control measures

- Is technically feasible

- Makes sense for the product strategy of the company

The evaluation is documented as part of the Change Evaluation List and results in a decision on whether the change is implemented and whether the Notified Body needs to be involved if the change is significant. These two decisions are documented in the change request activity.

If required, a Notified Body shall be informed before the implementation of changes. The multiple proposed changes can be batched and discussed in one session.

If the change request relates to the change of parts of the QMS, the change is assessed with the second part of the change evaluation list for “organizational changes”.

Participants

Product Manager

Head of Software Development

Regulatory Affairs Manager

| Input | Output |
| --- | --- |
| Change request | Completed change evaluation |

- Software Risk Management Process

Refer to RW-TF-FST-02-03(00) Risk Management

- Software Configuration Management Process

- Configuration management tools:

Configuration management is done via git, and GitHub is used for version identification and management.

- Configuration Management Items:

- Source and header files

- Project settings files

- Using library files

- Database configuration files

- Configuration Management Procedure:

- Software engineers create new branches that start with development.

- During development, software engineers continue intermediate commits in the branch they are working on.

- Once the implementation is complete, create a pull request.

- Other software engineers who didn’t write the code, and test engineer perform code reviews through the generated pull request.

- When the code review is complete, the software engineer merges the pull request.

- A new merge point is created at the development branch.

- Release Procedure:

The goal of each release is to uniquely identify each version and make all relevant files searchable.

- When a new software version is released, the software engineers create and tag the release through the branch.

- Tags contain version information, which is generated by complying with the criteria below.

- Version: A.B.C

| Category | Meaning |
| --- | --- |
| A | Major change |
| B | Minor change |
| C | Patch |

- Software Problem Resolution Process

- New Problem Evaluation

Redmine is used as a project management tool, and software engineers register new issues (bugs, defects, and needs improvement) as Activity within Redmine's Software project. Registered problems may arise from internal testing during the development process or from customer use, and the customer will share the problem through contact information of the person in charge if there is a problem during use.

For each problem report, the following must be entered:

- Affected medical device and version

- Severity classification (see below)

- Problem description incl. instructions to reproduce

We classify the severity of problems in the following categories:

| Severity Classification | Description |
| --- | --- |
| High | Causes new or changed risks to patients which are unacceptable. |
| Medium | May cause new or changed risks to patients which are acceptable. |
| Low | All other problems. |

For all problems classified as “Medium” or higher the person responsible for regulatory compliance (PRRC) must be informed who subsequently assesses it according to the SOP Vigilance.

Participants

Head of software development

Person responsible for regulatory compliance

| Input | Output |
| --- | --- |
| New problem | Problem report as Activity in Redmine |

- Root Cause Analysis and Procedure

For each registered activity, create a resolution plan and use Redmine's task distribution feature to assign the task to the appropriate team member. In addition, we can analyze and refer to whether similar problems have occurred in the past through Redmine's history, and in this case, we can link the activity through Redmine's Related Activity feature.

Participants

Head of software development

Software engineer

| Input | Output |
| --- | --- |
| Activity (Problem report) | Activity updated with cause and procedure |

- Implementation and Verification

Assigned team members follow established configuration management procedures to implement bug fixes in code. Once implemented, it is validated by other software engineers and test engineers, and the results are recorded in a Pull Request on GitHub. The assignee updates the activity with the content related to the activity and the generated pull request link and closes the activity by changing the status to 'Done'.

Head of software development

Software engineer

Test engineer

| Input | Output |
| --- | --- |
| Activity (Problem report) | Resolved/closed Activity |
| Updated source code |  |
| Closed pull request on GitHub |  |

- Documentation and Record Management

  - Software Life Cycle

Development proceeds according to the V model, and the output of each stage is traceable to the connected stage.

  - Activities at Each Software Life Cycle

| Activity | Output |
| --- | --- |
| Requirements capture | RW-TF-FST-02-04(00) Software Requirements Specifications (SRS) |
| Architecture design | RW-TF-FST-02-05(00) Software Architecture Description |
| Implementation | Source code |
| Verification & Validation | RW-TF-FST-02-08(00) Software V&V Report |

Attachment - Software – IEC 62304:2006+AMD1:2015

|  | Attachment - Software – IEC 62304:2006+AMD1:2015 |  | — |
| --- | --- | --- | --- |
| 4.3 | [A, B, C] Software safety classification |  | — |
|  | a) The manufacturer assigns to each software system a software safety class according to the risk of harm to the patient, operator, or other people resulting from a hazardous situation to which the software system can contribute in a worst-case-scenario | Refer to Clause 2 | A |
|  | The software system is software safety class A if: |  | — |
|  | – the software system not contribute to a hazardous situation; or |  |  |
|  | – the software system contribute to a hazardous situation which does not result in unacceptable risk after consideration of risk control measures external to the software system |  |  |
|  | The software system is software safety class B if: |  | — |
|  | – the software system contribute to a hazardous situation which results in unacceptable risk after consideration of risk control measures external to the software system and the resulting possible harm is non-serious injury |  |  |
|  | The software system is software safety class C if: |  | — |
|  | – the software system contribute to a hazardous situation which results in unacceptable risk after consideration of risk control measures external to the software system and the resulting possible harm is death or serious injury |  |  |
|  | For a software system initially classified as software safety class B or C, the manufacturer has implemented additional risk control measures external to the software system and subsequently has assigned a new software safety classification to the software system |  |  |
|  | c) The manufacturer documents the software safety class assigned to each software system in the risk management file | Refer to Clause 2 class : B | A |
|  | d) When a software system is decomposed into software items, and when a software item is decomposed into further software items, such software items inherit the software safety classification of the original software item (or software system) unless the manufacturer documents a rationale for classification into a different software safety class | N/A | N/A |
|  | A rationale explains how the new software items are segregated so that they may be classified separately | N/A | N/A |
|  | e) The manufacturer documents the software safety class of each software item if that class is different from the class of the software item from which it was created by decomposition | N/A | N/A |
|  | f) When applied to a group of software items, the manufacturer uses the processes and tasks which are required by the classification of the highest-classified software item in the group unless the manufacturer documents in the risk management file a rationale for using a lower classification | N/A | N/A |
|  | g) Class C requirements apply for each software system, until a software safety class is assigned | Refer to Clause 2 | A |
| 4.4 | [A, B, C] Legacy software |  | — |
|  | Clauses 5 through 9 have applied to demonstrate the compliance of legacy software | N/A | N/A |
|  | As alternative, clauses 4.4.2 through 4.4.5 have applied to demonstrate the compliance of legacy software | N/A | N/A |
| 4.4.2 | [A, B, C] Risk management activities |  | — |
|  | The manufacturer: Raywatt Inc. |  | A |
|  | a) assesses any feedback, including post-production information, on legacy software regarding incidents and / or near incidents, both from inside its own organization and / or from users | N/A | N/A |
|  | b) performs risk management activities associated with continued use of the legacy software | N/A | N/A |
|  | Considering the following aspects: |  |  |
|  | – integration of the legacy software in the overall medical device architecture | N/A | N/A |
|  | – continuing validity of risk control measures, implemented as part of the legacy software | N/A | N/A |
|  | – identification of hazardous situations associated with the continued use of the legacy software | N/A | N/A |
|  | – identification of potential causes of the legacy software contributing to a hazardous situations | N/A | N/A |
|  | – definition of risk control measures for each potential cause of the legacy software contributing to a hazardous situations | N/A | N/A |
| 4.4.3 | [A, B, C] Gap analysis |  |  |
|  | Based on the software safety class of the legacy software, the manufacturer performs a gap analysis of available deliverables against those required according to 5.2, 5.3, 5.7, and Clause 7 | N/A | N/A |
|  | a) The manufacturer assesses the continuing validity of available deliverables | N/A | N/A |
|  | b) Where gaps are identified, the manufacturer evaluates the potential reduction in risk resulting from the generation of the missing deliverables and associated activities | N/A | N/A |
|  | c) Based on this evaluation, the manufacturer determines the deliverables to be created and associated activities to be performed | N/A | N/A |
|  | Software system test records are the minimum deliverables to be created | N/A | N/A |
| 4.4.4 | [A, B, C] Gap closure activities |  |  |
|  | a) The manufacturer establishes and executes a plan to generate the identified deliverables | N/A | N/A |
|  | Objective evidences have used to generate required deliverables without performing activities required by 5.2, 5.3, 5.7 and Clause 7 | N/A | N/A |
|  | b) The plan addresses the use of the problem resolution process for handling problems detected in the legacy software and deliverables in accordance with Clause 9 | N/A | N/A |
|  | c) Changes to the legacy software have performed in accordance with Clause 6. | N/A | N/A |
| 4.4.5 | [A, B, C] Rationale for use of legacy software |  |  |
|  | The manufacturer documents the version of the legacy software together with a rationale for the continued use of the legacy software | N/A | N/A |

| 5 | SOFTWARE DEVELOPMENT PROCESS |  | — |
| --- | --- | --- | --- |
| 5.1 | Software development planning |  | — |
| 5.1.1 | [A, B, C] The manufacturer establishes a software development plan (or plans) for conducting the activities of the software development process appropriate to the scope, magnitude, and software safety classifications of the software system to be developed. | Refer to Clause 3 | A |
|  | The software development life cycle model is either fully defined or be referenced in the plan (or plans). | Refer to Clause 3 | A |
|  | The plan addresses the following: |  |  |
|  | a) the processes to be used in the development of the software system | Refer to Clause 3 | A |
|  | b) the deliverables (includes documentation) of the activities and tasks | Refer to Clause 8 | A |
|  | c) traceability between system requirements, software requirements, software system test, and risk control measures implemented in software | Refer to Clause 3 | A |
|  | d) software configuration and change management, including soup configuration items and software used to support development | Refer to RW-TF-FST-02-02(00) Software Description | A |
|  | e) software problem resolution for handling problems detected in the medical device software, deliverables and activities at each stage of the life cycle | Refer to Clause 3 | A |
| 5.1.2 | [A, B, C] The manufacturer updates the plan, as appropriate, as development proceeds | Refer to Clause 3 | A |
| 5.1.3 | [A, B, C] Software development plan reference to system design and development |  |  |
|  | a) As inputs for software development, system requirements are referenced in the software development plan by the manufacturer | Refer to Clause 3 | A |
|  | b) In the software development plan, the manufacturer includes or references procedures for coordinating the software development with the system development necessary to satisfy 4.1 (such as system integration, verification, and validation) | Refer to Clause 3 | A |
| 5.1.4 | [C] Associated with the development of software items of class C, in the software development plan are included or referenced: |  |  |
|  | a) standards | N/A | N/A |
|  | b) methods | N/A | N/A |
|  | c) tools | N/A | N/A |
| 5.1.5 | [B, C] The manufacturer includes or references in the software development plan, a plan to integrate the software items (including soup) and performs testing during integration | Refer to RW-TF-FST-02-02(00) Software Description | A |
| 5.1.6 | [A, B, C] In the software development plan, the following verification information are included or referenced: |  |  |
|  | a) deliverables requiring verification | Refer to Clause 3 | A |
|  | b) the required verification tasks for each life cycle activity | Refer to Clause 3 | A |
|  | c) milestones at which the deliverables are verified | Refer to Clause 3 | A |
|  | d) the acceptance criteria for verification of the deliverables | Refer to Clause 3 | A |
| 5.1.7 | [A, B, C] In the software development plan the manufacturer includes or references a plan to conduct the activities and tasks of the software risk management process, including the management of risks relating to soup | Refer to Clause 5 | A |
| 5.1.8 | [A, B, C] In the software development plan the manufacturer includes or references information about the documents to be produced during the software development life cycle | Refer to Clause 8 | A |
|  | For each identified document or type of document the following information has included or referenced: |  |  |
|  | a) title, name or naming convention | Each document | A |
|  | b) purpose | Each document | A |
|  | c) procedures and responsibilities for development, review, approval and modification | Each document | A |
| 5.1.9 | [A, B, C] The manufacturer includes or references software configuration management information in the software development plan | Refer to Clause 6 | A |
|  | The software configuration management information includes or references: |  |  |
|  | a) the classes, types, categories or lists of items to be controlled | Refer to Clause 6 | A |
|  | b) the software configuration management activities and tasks | Refer to Clause 6 | A |
|  | c) the organization(s) responsible for performing software configuration management and activities | Refer to Clause 6 | A |
|  | d) their relationship with other organizations, such as software development or maintenance | Refer to Clause 6 | A |
|  | e) when the items are to be placed under configuration control | Refer to Clause 6 | A |
|  | f) when the problem resolution process is to be used | Refer to Clause 6 | A |
| 5.1.10 | [B, C] The items to be controlled include tools, items or settings, used to develop the medical device software, which could impact the medical device software | Refer to Clause 6 | A |
| 5.1.11 | [B, C] The manufacturer plans to place configuration items under documented configuration management control before they are verified | Refer to Clause 6 | A |
| 5.1.12 | [B, C] In the software development plan the manufacturer includes or references a procedure for: |  |  |
|  | a) identifying categories of defects that may be introduced based on the selected programming technology that are relevant to their software system | Refer to Clause 7 | A |
|  | b) documenting evidence that demonstrates that these defects do not contribute to unacceptable risk | Refer to Clause 7 | A |
| 5.2 | Software requirements analysis |  | — |
| 5.2.1 | [A, B, C] For each software system of the medical device, the manufacturer defines and documents software system requirements from the system level requirements | Refer to Clause 3 | A |
| 5.2.2 | [A, B, C] As appropriate to the medical device software, the manufacturer includes in the software requirements: |  |  |
|  | a) functional and capability requirements | Refer to Clause 3 | A |
|  | b) software system inputs and outputs | Refer to Clause 3 | A |
|  | c) interfaces between the software system and other systems | Refer to Clause 3 | A |
|  | d) software-driven alarms, warnings, and operator messages | Refer to Clause 3 | A |
|  | e) security requirements | Refer to Clause 3 | A |
|  | f) user interface requirements implemented by software | Refer to Clause 3 | A |
|  | g) data definition and database requirements | Refer to Clause 3 | A |
|  | h) installation and acceptance requirements of the delivered medical device software at the operation and maintenance site or sites | Refer to Clause 3 | A |
|  | i) requirements related to methods of operation and maintenance | Refer to Clause 3 | A |
|  | j) requirements related to IT-network aspects | N/A | N/A |
|  | k) user maintenance requirements | Refer to Clause 3 | A |
|  | l) regulatory requirements | Refer to Clause 3 | A |
| 5.2.3 | [B, C] The manufacturer includes risk control measures implemented in software in the requirements as appropriate to the medical device software | Refer to Clause 5 | A |
| 5.2.4 | [A, B, C] The manufacturer re-evaluates the medical device risk analysis when software requirements are established and update it as appropriate | Refer to Clause 5 | A |
| 5.2.5 | [A, B, C] The manufacturer ensures that existing requirements, including system requirements, are re-evaluated and updated as appropriate as a result of the software requirements analysis activity | Refer to Clause 5 | A |
| 5.2.6 | [A, B, C] The manufacturer verifies and documents that the software requirements: |  |  |
|  | a) implement system requirements including those relating to risk control | Refer to Clause 5 | A |
|  | b) do not contradict one another | Refer to Clause 5 | A |
|  | c) are expressed in terms that avoid ambiguity | Refer to Clause 5 | A |
|  | d) are stated in terms that permit establishment of test criteria and performance of tests | Refer to Clause 5 | A |
|  | e) can be uniquely identified | Refer to Clause 5 | A |
|  | f) are traceable to system requirements or other source | Refer to Clause 5 | A |
| 5.3 | Software architectural design |  |  |
| 5.3.1 | [B, C] The manufacturer transforms the requirements for the medical device software into a documented architecture that describes the software’s structure and identifies the software items | Refer to Clause 3 | A |
| 5.3.2 | [B, C] The manufacturer develops and documents an architecture for the interfaces between the software items and the components external to the software items (both software and hardware), and between the software items | Refer to Clause 3 | A |
| 5.3.3 | [B, C] If a software item is identified as soup, the manufacturer specifies functional and performance requirements for the soup item that are necessary for its intended use | Refer to Clause 3 | A |
| 5.3.4 | [B, C] If a software item is identified as soup, the manufacturer specifies the system hardware and software necessary to support the proper operation of the soup item | Refer to Clause 3 | A |
| 5.3.5 | [C] The manufacturer identifies any segregation between software items that is necessary for risk control, and states how to ensure that such segregation is effective | N/A | N/A |
| 5.3.6 | [B, C] The manufacturer verifies and documents that: |  |  |
|  | a) the architecture of the software implements system and software requirements including those relating to risk control | Refer to Clause 3 | A |
|  | b) the software architecture is able to support interfaces between software items and between software items and hardware | Refer to Clause 3 | A |
|  | c) the medical device architecture supports proper operation of any soup items | Refer to Clause 3 | A |
| 5.4 | Software detailed design |  |  |
| 5.4.1 | [B, C] The manufacturer subdivides the software until it is represented by software units | Refer to Clause 3 | A |
| 5.4.2 | [C] The manufacturer documents a design with enough detail to allow correct implementation of each software unit | N/A | N/A |
| 5.4.3 | [C] The manufacturer documents a design for any interfaces between the software unit and external components (hardware or software), as well as any interfaces between software units, detailed enough to implement each software unit and its interfaces correctly | N/A | N/A |
| 5.4.4 | [C] The manufacturer verifies and documents that the software detailed design: |  |  |
|  | a) implements the software architecture | N/A | N/A |
|  | b) is free from contradiction with the software architecture | N/A | N/A |
| 5.5 | Software unit implementation |  |  |
| 5.5.1 | [A, B, C] The manufacturer implements each software unit | Refer to Clause 3 | A |
| 5.5.2 | [B, C] The manufacturer establishes strategies, methods and procedures for verifying the software units | Refer to Clause 3 | A |
|  | Where verification is done by testing, the test procedures are evaluated for adequacy | Refer to Clause 3 | A |
| 5.5.3 | [B, C] The manufacturer establishes acceptance criteria for software units prior to integration into larger software items as appropriate, and ensures that software units meet acceptance criteria | Refer to Clause 3 | A |
| 5.5.4 | [C] When present in the design, the manufacturer includes additional acceptance criteria as appropriate for: |  |  |
|  | a) proper event sequence | N/A | N/A |
|  | b) data and control flow | N/A | N/A |
|  | c) planned resource allocation | N/A | N/A |
|  | d) fault handling (error definition, isolation, and recovery) | N/A | N/A |
|  | e) initialisation of variables | N/A | N/A |
|  | f) self-diagnostics | N/A | N/A |
|  | g) memory management and memory overflows | N/A | N/A |
|  | h) boundary conditions | N/A | N/A |
| 5.5.5 | [B, C] The manufacturer performs the software unit verification and documents the results | Refer to Clause 3 | A |
| 5.6 | Software integration and integration testing |  |  |
| 5.6.1 | [B, C] The manufacturer integrates the software units in accordance with the integration plan | Refer to Clause 3 | A |
| 5.6.2 | [B, C] The manufacturer verifies that the software units have been integrated into software items and/or the software system in accordance with the integration plan and retains records of the evidence of such verification | Refer to Clause 3 | A |
| 5.6.3 | [B, C] The manufacturer tests the integrated software items in accordance with the integration plan and documents the results | Refer to Clause 3 | A |
| 5.6.4 | [B, C] For software integration testing, the manufacturer addresses whether the integrated software item performs as intended | Refer to Clause 3 | A |
| 5.6.5 | [B, C] The manufacturer evaluates the integration test procedures for adequacy | Refer to Clause 3 | A |
| 5.6.6 | [B, C] When software items are integrated, the manufacturer conducts regression testing appropriate to demonstrate that defects have not been introduced into previously integrated software | Refer to Clause 3 | A |
| 5.6.7 | [B, C] The manufacturer: |  |  |
|  | a) documents the test result (pass/fail and a list of anomalies) | Refer to Clause 3 | A |
|  | b) retains sufficient records to permit the test to be repeated | Refer to Clause 3 | A |
|  | c) identifies the tester | Refer to Clause 3 | A |
| 5.6.8 | [B, C] The manufacturer enters anomalies found during software integration and integration testing into a software problem resolution process | Refer to Clause 3 | A |
| 5.7 | Software system testing |  |  |
| 5.7.1 | [A, B, C] Establish tests for software requirements |  | — |
|  | a) The manufacturer establishes and performs a set of tests, expressed as input stimuli, expected outcomes, pass/fail criteria and procedures, for conducting software system testing, such that all software requirements are covered | Refer to Clause 3 | A |
|  | b) The manufacturer evaluates the adequacy of verification strategies and test procedures. | Refer to Clause 3 | A |
| 5.7.2 | [A, B, C] The manufacturer enters anomalies found during software system testing into a software problem resolution process | Refer to Clause 3 | A |
| 5.7.3 | [A, B, C] When changes are made during software system testing, the manufacturer: |  |  |
|  | a) repeats tests, performs modified tests or performs additional tests, as appropriate, to verify the effectiveness of the change in correcting the problem | Refer to Clause 3 | A |
|  | b) conducts testing appropriate to demonstrate that unintended side effects have not been introduced | Refer to Clause 3 | A |
|  | c) performs relevant risk management activities as defined in 7.4 | Refer to Clause 3 | A |
| 5.7.4 | [A, B, C] Evaluate software system testing |  |  |
|  | The manufacturer evaluates the appropriateness of verification strategies and test procedures | Refer to Clause 3 | A |
|  | The manufacturer verifies that: |  |  |
|  | a) all software requirements have been tested or otherwise verified | Refer to Clause 3 | A |
|  | b) the traceability between software requirements and tests or other verification is recorded | Refer to Clause 3 | A |
|  | c) test results meet the required pass/fail criteria | Refer to Clause 3 | A |
| 5.7.5. | [A, B, C] In order to support the repeatability of tests, the manufacturer documents: |  |  |
|  | a) a reference to test case procedures showing required actions and expected results | Refer to Clause 3 | A |
|  | b) the test result (pass/fail and a list of anomalies) | Refer to Clause 3 | A |
|  | c) the version of software tested | Refer to Clause 3 | A |
|  | d) relevant hardware and software test configurations | Refer to Clause 3 | A |
|  | e) relevant test tools | Refer to Clause 3 | A |
|  | f) date tested | Refer to Clause 3 | A |
|  | g) the identity of the person responsible for executing the test and recording the test results | Refer to Clause 3 | A |
| 5.8 | Software release for utilization at a system level |  |  |
| 5.8.1 | [A, B, C] The manufacturer ensures that all software verification activities has been completed and the results evaluated before the software is released | Refer to Clause 3 | A |
| 5.8.2 | [A, B, C] The manufacturer documents all known residual anomalies | Refer to Clause 3 | A |
| 5.8.3 | [B, C] The manufacturer ensured that all known residual anomalies have been evaluated to ensure that they do not contribute to an unacceptable risk | Refer to Clause 3 | A |
| 5.8.4 | [A, B, C] The manufacturer documented the version of the medical device software that is being released | Refer to Clause 3 | A |
| 5.8.5 | [B, C] The manufacturer documents the procedure and environment used to create the released software | Refer to Clause 3 | A |
| 5.8.6 | [B, C] The manufacturer ensures that all software development plan (or maintenance plan) activities and tasks are complete along with the associated documentation | Refer to Clause 3 | A |
| 5.8.7 | [A, B, C] For at least a period of time determined as the longer of: the life time of the medical device software as defined by the manufacturer or a time specified by relevant regulatory requirements, the manufacturer archives: |  |  |
|  | a) the medical device software and configuration items | Refer to Clause 3 | A |
|  | b) the documentation | Refer to Clause 3 | A |
| 5.8.8 | [A, B, C] The manufacturer establishes procedures to ensure that the released medical device software can be reliably delivered to the point of use without corruption or unauthorised change | Refer to Clause 3 | A |
|  | These procedures address the production and handling of media containing the medical device software including as appropriate: |  |  |
|  | – replication | N/A | N/A |
|  | – media labelling | N/A | N/A |
|  | – packaging | Refer to User Manual | A |
|  | – protection | Refer to User Manual | A |
|  | – storage | Refer to User Manual | A |
|  | – delivery | Refer to User Manual | A |

| 7 | SOFTWARE RISK MANAGEMENT PROCESS |  | — |
| --- | --- | --- | --- |
| 7.1 | Analysis of software contributing to hazardous situations |  | — |
| 7.1.1 | [B, C] The manufacturer identifies software items that could contribute to a hazardous situation identified in the medical device risk analysis activity of ISO 14971 | Refer to Clause 5 | A |
| 7.1.2 | [B, C] The manufacturer identifies potential causes of the software item identified above contributing to a hazardous situation | Refer to Clause 5 | A |
|  | The manufacturer considers potential causes including, as appropriate: |  |  |
|  | a) incorrect or incomplete specification of functionality | Refer to Clause 5 | A |
|  | b) software defects in the identified software item functionality | Refer to Clause 5 | A |
|  | c) failure or unexpected results from soup | Refer to Clause 5 | A |
|  | d) hardware failures or other software defects that could result in unpredictable software operation | Refer to Clause 5 | A |
|  | e) reasonably foreseeable misuse | Refer to Clause 5 | A |
| 7.1.3 | [B, C] If failure or unexpected results from soup is a potential cause of the software item contributing to a hazardous situation, the manufacturer evaluates as a minimum any anomaly list published by the supplier of the soup item relevant to the version of the soup item used in the medical device to determine if any of the known anomalies result in a sequence of events that could result in a hazardous situation | Refer to Clause 5 | A |
| 7.1.4 | [B, C] The manufacturer documents in the risk management file potential causes of the software item contributing to a hazardous situation | Refer to Clause 5 | A |
| 7.2 | Risk control measures |  | — |
| 7.2.1 | [B, C] For each case documented in the risk management file where a software item could contribute to a hazardous situation, the manufacturer defines and documents risk control measures in accordance with ISO 14971 | Refer to Clause 5 | A |
| 7.2.2 | [B, C] If a risk control measure is implemented as part of the functions of a software item, the manufacturer: |  |  |
|  | a) includes the risk control measure in the software requirements | Refer to Clause 5 | A |
|  | b) assigns to each software item that contributes to the implementation of a risk control measure a software safety class based on the risk that the risk control measure is controlling | Refer to Clause 5 | A |
|  | c) develops the software item in accordance with Clause 5 | Refer to Clause 5 | A |
| 7.3 | Verification of risk control measures |  | — |
| 7.3.1 | [B, C] The implementation of each risk control measure documented in 7.2 is verified, and this verification is documented | Refer to Clause 5 | A |
|  | The manufacturer reviewers the risk control measure and determines if it could result in a new hazardous situation | Refer to Clause 5 | A |
| 7.3.3 | [B, C] The manufacturer documents traceability of software hazards as appropriate: |  |  |
|  | a) from the hazardous situation to the software item | Refer to Clause 5 | A |
|  | b) from the software item to the specific software cause | Refer to Clause 5 | A |
|  | c) from the software cause to the risk control measure | Refer to Clause 5 | A |
|  | d) from the risk control measure to the verification of the risk control measure | Refer to Clause 5 | A |
| 7.4 | Risk management of software changes |  | — |
| 7.4.1 | [A, B, C] The manufacturer analyses changes to the medical device software (including soup) to determine whether: |  |  |
|  | a) additional potential causes are introduced contributing to a hazardous situation | Refer to Clause 5 | A |
|  | b) additional software risk control measures are required | Refer to Clause 5 | A |
| 7.4.2 | [B, C] The manufacturer analyses changes to the software, including changes to soup, to determine whether the software modification could interfere with existing risk control measures | Refer to Clause 5 | A |
| 7.4.3 | [B, C] The manufacturer performs relevant risk management activities defined in 7.1, 7.2 and 7.3 based on these analyses | Refer to Clause 5 | A |

| 8 | SOFTWARE CONFIGURATION MANAGEMENT PROCESS |  | — |
| --- | --- | --- | --- |
| 8.1 | Configuration identification |  | — |
| 8.1.1 | [A, B, C] The manufacturer establishes a scheme for the unique identification of configuration items and their versions to be controlled according to the development and configuration planning specified in 5.1 | Refer to Clause 6 | A |
| 8.1.2 | [A, B, C] For each soup configuration item being used, including standard libraries, the manufacturer documents: |  |  |
|  | a) the title | Refer to Clause 6 | A |
|  | b) the manufacturer | Refer to Clause 6 | A |
|  | c) the unique soup designator | Refer to Clause 6 | A |
| 8.1.3 | [A, B, C] The manufacturer documents the set of configuration items and their versions that comprise the software system configuration | Refer to Clause 6 | A |
| 8.2 | Change control |  | — |
| 8.2.1 | [A, B, C] The manufacturer changes configuration items identified to be controlled according to 8.1 only in response to an approved change request | Refer to Clause 4 | A |
| 8.2.2 | [A, B, C] The manufacturer implements the change as specified in the change request | Refer to Clause 4 | A |
|  | The manufacturer identifies and performs any activity that needs to be repeated as a result of the change, including changes to the software safety classification of software systems and software items | Refer to Clause 4 | A |
| 8.2.3 | [A, B, C] The manufacturer verifies the change, including repeating any verification that has been invalidated by the change and taking into account 5.7.3 and 9.7 | Refer to Clause 4 | A |
| 8.2.4 | [A, B, C] The manufacturer maintains records of the relationships and dependencies between: |  |  |
|  | a) change request | Refer to Clause 4 | A |
|  | b) relevant problem report | Refer to Clause 4 | A |
|  | c) approval of the change request | Refer to Clause 4 | A |
| 8.3 | [A, B, C] The manufacturer retains retrievable records of the history of controlled configuration items including system configuration | Refer to Clause 4 | A |

| 9 | SOFTWARE PROBLEM RESOLUTION PROCESS |  | — |
| --- | --- | --- | --- |
| 9.1 | [A, B, C] The manufacturer prepares a problem report for each problem detected in the medical device software | Refer to Clause 7 | A |
|  | Problem reports include a statement of criticality (for example, effect on performance, safety, or security) as well as other information that may aid in the resolution of the problem (for example, devices affected, supported accessories affected) | Refer to Clause 7 | A |
| 9.2 | [A, B, C] The manufacturer: |  |  |
|  | a) investigates the problem and if possible identify the causes | Refer to Clause 7 | A |
|  | b) evaluates the problem’s relevance to safety using the software risk management process | Refer to Clause 7 | A |
|  | c) documents the outcome of the investigation and evaluation | Refer to Clause 7 | A |
|  | d) creates a change request(s) for actions needed to correct the problem, or document the rationale for taking no action | Refer to Clause 7 | A |
| 9.3 | [A, B, C] The manufacturer advises relevant parties of the existence of the problem, as appropriate | Refer to Clause 7 | A |
| 9.4 | [A, B, C] The manufacturer approves and implements all change requests, observing the requirements of the change control process | Refer to Clause 7 | A |
| 9.5 | [A, B, C] The manufacturer maintains records of problem reports and their resolution including their verification | Refer to Clause 7 | A |
|  | The manufacturer updates the risk management file as appropriate | Refer to Clause 7 | A |
| 9.6 | [A, B, C] The manufacturer performs analysis to detect trends in problem reports | Refer to Clause 7 | A |
| 9.7 | [A, B, C] The manufacturer verifies resolutions to determine whether: |  |  |
|  | a) problem has been resolved and the problem report has been closed | Refer to Clause 7 | A |
|  | b) adverse trends have been reversed | Refer to Clause 7 | A |
|  | c) change requests have been implemented in the appropriate medical device software and activities | Refer to Clause 7 | A |
|  | d) additional problems have been introduced | Refer to Clause 7 | A |
| 9.8 | [A, B, C] When testing, retesting or regression testing software items and systems following a change, the manufacturer includes in the test documentation: |  |  |
|  | a) test results | Refer to Clause 3 | A |
|  | b) anomalies found | Refer to Clause 3 | A |
|  | c) the version of software tested | Refer to Clause 3 | A |
|  | d) relevant hardware and software test configurations | Refer to Clause 3 | A |
|  | e) relevant test tools | Refer to Clause 3 | A |
|  | f) date tested | Refer to Clause 3 | A |
|  | g) identification of the tester | Refer to Clause 3 | A |