## Integration Test
SIT ID | Item | Verification | Expected Result | Actual Result | Pass/Fail
[5 empty] | Test 1 | Test 2 | Test 3
SIT-01 | Page display | Check if the Patient List screen is displayed after loading | Patient List screen is displayed. | Patient List screen is displayed. | Pass | Pass | Pass
| Dialog display | Check if the Import Dialog (Popup) appears when clicking the Import button | Import Popup is displayed. | Import Popup is displayed. | Pass | Pass | Pass
SIT-02 | Data input into Database | Check if a new patient is added via New Patient | A new patient is added. | A new patient is added. | Pass | Pass | Pass
| Data retrieval from Database | Check if the patient list is displayed on the Patient List screen | The patient list is displayed on the Patient List screen. | The patient list is displayed on the Patient List screen. | Pass | Pass | Pass
| Data modification in Database | Check if patient information is modified in Edit Patient | Patient information is modified. | Patient information is modified. | Pass | Pass | Pass
| Data deletion in Database | Check if the patient is removed from the list after deletion | Patient information is deleted. | Patient information is deleted. | Pass | Pass | Pass
SIT-03 | BLDC Motor rotation control | Check if the Catheter connector rotates when connected to RJ | A red light is displayed at the tip of the Catheter. | A red light is displayed at the tip of the Catheter. | Pass | Pass | Pass
| Pullback Motor control to adjust Linear Stage | Check if the Catheter connector moves when connected to RJ | A red light is displayed at the tip of the Catheter. | A red light is displayed at the tip of the Catheter. | Pass | Pass | Pass
| Acquisition of Catheter connection status | Check if VLD turns on at the tip of the Catheter | A red light is displayed at the tip of the Catheter. | A red light is displayed at the tip of the Catheter. | Pass | Pass | Pass
SIT-04 | Visible Light power control | Check VLD at the tip of the Catheter | A red light is displayed at the tip of the Catheter. | A red light is displayed at the tip of the Catheter. | Pass | Pass | Pass
| Laser power control | Check if real-time OCT images are displayed on the LiveView screen | OCT image is displayed. | OCT image is displayed. | Pass | Pass | Pass
| Delay Line Motor control to adjust Linear Stage | Check if Calibration works on the LiveView screen | Calibration is activated. | Calibration is activated. | Pass | Pass | Pass
| Polarization Motor control for angle adjustment | Check if Calibration works on the LiveView screen | Calibration is activated. | Calibration is activated. | Pass | Pass | Pass
SIT-05 | Acquisition of OCT Signal from Acquisition Board | Check if real-time OCT images are displayed on the LiveView screen | OCT image is displayed. | OCT image is displayed. | Pass | Pass | Pass
| Conversion of OCT Signal to Image format | Check if real-time OCT images are displayed on the LiveView screen | OCT image is displayed. | OCT image is displayed. | Pass | Pass | Pass
| Loading Raw Image file | Check if OCT images are displayed in Review | OCT image is displayed. | OCT image is displayed. | Pass | Pass | Pass
| Saving Raw Image file | Open a saved Case to check the OCT image | OCT image is displayed. | OCT image is displayed. | Pass | Pass | Pass
| Construction of L-View Image from Cross-Section Image | Check L-View in Review | L-View is displayed. | L-View is displayed. | Pass | Pass | Pass
| Degree adjustment of OCT Image to rotate L-View Image | Check if L-View rotates when adjusting the Degree Indicator in Review | L-View rotates while being displayed. | L-View rotates while being displayed. | Pass | Pass | Pass
| Frame position adjustment to modify Cross-Section Image | Check if Cross-Section Image changes when adjusting the Frame Indicator in Review | Cross-Section image is updated. | Cross-Section image is updated. | Pass | Pass | Pass
| OCT Image Brightness adjustment | Check brightness adjustment in Review | Brightness of the image is adjusted. | Brightness of the image is adjusted. | Pass | Pass | Pass
| OCT Image Contrast adjustment | Check contrast adjustment in Review | Contrast of the image is adjusted. | Contrast of the image is adjusted. | Pass | Pass | Pass
SIT-06 | Extraction of Lumen Contour information from OCT Image | Check Lumen Contour information after Recording | Lumen Contour is displayed in the Cross-Section. | Lumen Contour is displayed in the Cross-Section. | Pass | Pass | Pass
| Extraction of Side Branch information from OCT Image | Check Side Branch information after Recording | Side Branch location is displayed in the Lumen Profile. | Side Branch location is displayed in the Lumen Profile. | Pass | Pass | Pass
| Extraction of Stent information from OCT Image | Check Stent information after Recording | Stent location is displayed in the Cross-Section. | Stent location is displayed in the Cross-Section. | Pass | Pass | Pass
| Extraction of Guidewire information from OCT Image | Check Guidewire information after Recording (3D screen) | Guidewire is displayed in the Cut-View of the 3D image. | Guidewire is displayed in the Cut-View of the 3D image. | Pass | Pass | Pass
| Extraction of Calcium information from OCT Image | Check Calcium information after Recording | Calcium location is displayed in the Cross-Section. | Calcium location is displayed in the Cross-Section. | Pass | Pass | Pass
SIT-07 | 3D Cut-View image rendering | Check Cut-View images in the 3D screen | 3D Object is displayed in the Cut-View area. | 3D Object is displayed in the Cut-View area. | Pass | Pass | Pass
| 3D Fly-Through image rendering | Check Fly-Through images in the 3D screen | 3D Object is displayed in the Fly-Through area. | 3D Object is displayed in the Fly-Through area. | Pass | Pass | Pass
| 3D Object rotation | Check 3D Object rotation when adjusting the Degree Indicator in the 3D screen | 3D Object rotates while being displayed. | 3D Object rotates while being displayed. | Pass | Pass | Pass
| Camera Object position adjustment | Check 3D Object position adjustment when manipulating the Frame Indicator in the 3D screen | Position of the 3D Object is updated. | Position of the 3D Object is updated. | Pass | Pass | Pass
| Modification of Tissue/Lumen/Guidewire/Side Branch rendering method | Check rendering changes when modifying 3D screen options | 3D Object display is updated based on the selected Option. | 3D Object display is updated based on the selected Option. | Pass | Pass | Pass
SIT-08 | Measurement of Area, Diameter, %AS, %DS | Check Lumen information (Area, Diameter, %AS, %DS) in Review | Lumen information (Area, Diameter, %AS, %DS) is displayed. | Lumen information (Area, Diameter, %AS, %DS) is displayed. | Pass | Pass | Pass
| Measurement of Expansion, Apposition | Check Stent information (Expansion, Apposition) in Review | Stent information (Expansion, Apposition) is displayed. | Stent information (Expansion, Apposition) is displayed. | Pass | Pass | Pass
| Measurement of Area and Diameter of a specific region | Check Area and Diameter by drawing a region in Review | An area is added to the image, and Area, Length information is displayed. | An area is added to the image, and Area, Length information is displayed. | Pass | Pass | Pass
| Measurement of Length between two points | Check Length by drawing a line in Review | A line is added to the image, and Length information is displayed. | A line is added to the image, and Length information is displayed. | Pass | Pass | Pass
| Annotation creation | Check annotation display in Review | Annotations are added to the image. | Annotations are added to the image. | Pass | Pass | Pass
SIT-09 | Angio Server process control | Check if the Angio icon is activated in the bottom left | Angio icon is activated in the bottom left corner. | Angio icon is activated in the bottom left corner. | Pass | Pass | Pass
| Receiving Angio Image from Angio Server | Check Angio images in the Recording screen | Real-time Angio video is displayed. | Real-time Angio video is displayed. | Pass | Pass | Pass
| Loading Angio Raw Image file | Check if Angio images are displayed in Review | Angio video is displayed. | Angio video is displayed. | Pass | Pass | Pass
| Saving Angio Raw Image file | Open a saved Case to check the Angio image | Angio video is displayed. | Angio video is displayed. | Pass | Pass | Pass
| Acquiring Angio Image from Angio equipment | Check Angio images in the Recording screen | Real-time Angio video is displayed. | Real-time Angio video is displayed. | Pass | Pass | Pass
| Transmitting acquired Angio Image to Angio Client | Check Angio images in the Recording screen | Real-time Angio video is displayed. | Real-time Angio video is displayed. | Pass | Pass | Pass
SIT-10 | Generating Lumen Profile Image using Lumen Contour information | Check Lumen Profile in Review | Lumen Profile is displayed. | Lumen Profile is displayed. | Pass | Pass | Pass
| Generating Review screen as an Export Image | Check image results in Export | The image is displayed according to the selected Export Option. | The image is displayed according to the selected Export Option. | Pass | Pass | Pass
| Converting Export Image to Standard format (MP4, JPEG, Bitmap, TIFF) | Check saved format in Export results | Exported files are saved in the corresponding format. | Exported files are saved in the corresponding format. | Pass | Pass | Pass
| Converting Export Image to DICOM format | Check stored DICOM type in Export results | Exported files are saved in DICOM format. | Exported files are saved in DICOM format. | Pass | Pass | Pass
SIT-11 | Encrypting important information during data extraction | Check if file contents are encrypted in Export results | Exported file contents are encrypted. | Exported file contents are encrypted. | Pass | Pass | Pass
SIT-12 | Logging functionality for access, query, creation, modification, and deletion | Check Log file content by retrieving Log files from the Setting menu | Log files record access, query, creation, modification, and deletion history. | Log files record access, query, creation, modification, and deletion history. | Pass | Pass | Pass
SIT-13 | Only authorized users can access the system by entering ID/PW | Check if system access requires user ID/PW entry | User must enter the correct ID/PW to access. | User must enter the correct ID/PW to access. | Pass | Pass | Pass
SIT-14 | Only USB data communication is allowed, with wired and wireless communication restricted | Check if the device has a LAN port and supports Wi-Fi functionality | The device has no LAN port, and Wi-Fi functionality is restricted. | The device has no LAN port, and Wi-Fi functionality is restricted. | Pass | Pass | Pass