## System Test
SST ID | Item | Test Procedure | Expected Result | Actual Result | Pass/Fail
[5 empty] | Test 1 | Test 2 | Test 3
SST-01 | UI Display | 1. Access Home (Patient List) Screen 2. Execute Import Function | The main screen and pop-ups are displayed. | The main screen and pop-ups are displayed. | Pass | Pass | Pass
SST-02 | DB Processing | 1. Add new patient information in the Add Patient screen 2. Modify patient information in the Edit Patient screen 3. Delete patient information in the Edit Patient screen 4. Check patient information on the Patient List screen | Patient information can be added, modified, deleted, and viewed. | Patient information can be added, modified, deleted, and viewed. | Pass | Pass | Pass
SST-03 | Catheter Connection | 1. Configure recording settings on the Recording Settings screen and proceed to Next 2. Connect RJ to Catheter 3. Press the Live button on the Live View screen 4. Execute Auto Calibration on the Calibration screen | OCT images are displayed, and OCT image Calibration is executed. | OCT images are displayed, and OCT image Calibration is executed. | Pass | Pass | Pass
SST-04 | Recording & Review | 1. Recording 1-1. Start Recording in the Live View screen 1-2. Execute Ready and Start in the Recording screen 1-3. Select Accept and Continue on the Image Quality Check screen 2. Review (2D) 2-1. Activate OCT mode in the Review (2D) screen 2-2. Activate OCT + Angio mode in the Review (2D) screen 2-3. Activate Measure-Auto in the Review (2D) screen 2-4. Activate Measure-Manual and perform measurement in the Review (2D) screen 2-5. Operate Presentation in the Review (2D) screen 2-6. Activate Lumen Profile in the Review (2D) screen 2-7. Activate L-View in the Review (2D) screen 2-8. Adjust Frame/Degree Indicator in the Review (2D) screen 3. Review (3D) 3-1. Change 3D Option in the Review (3D) screen 3-2. Adjust Frame/Degree Indicator in the Review (3D) screen (Note: The displayed content may vary depending on Pre/Post/Follow-up cases.) | 1. Recording Screen - OCT and Angio images are recorded. 2. Review (2D) Screen - Cross-Section, Angio, Lumen Profile, and L-View are displayed. - Lumen Contour, Side Branch, Stent, and Calcium information are displayed. - Automatically measured values such as Area, Diameter, %AS, %DS, Expansion, and Apposition are displayed. - Manually measured values such as Area, Diameter, Length, and Annotations are displayed. - Cross-Section and L-View are configured according to Indicator and Presentation. 3. Review (3D) Screen - 3D Objects are displayed in Cut-View and Fly-Through formats. - Objects are selectively displayed based on option selection. - Cut-View and Fly-Through are configured according to the Indicator. | 1. Recording Screen - OCT and Angio images are recorded. 2. Review (2D) Screen - Cross-Section, Angio, Lumen Profile, and L-View are displayed. - Lumen Contour, Side Branch, Stent, and Calcium information are displayed. - Automatically measured values such as Area, Diameter, %AS, %DS, Expansion, and Apposition are displayed. - Manually measured values such as Area, Diameter, Length, and Annotations are displayed. - Cross-Section and L-View are configured according to Indicator and Presentation. 3. Review (3D) Screen - 3D Objects are displayed in Cut-View and Fly-Through formats. - Objects are selectively displayed based on option selection. - Cut-View and Fly-Through are configured according to the Indicator. | Pass | Pass | Pass
SST-05 | Export | 1. Execute Export Function 2. Select Export Type and the Patient Case to Extract 3. Choose Options according to Each Export Type and Execute Export | Images are generated according to the selected Export Type and Option settings. | Images are generated according to the selected Export Type and Option settings. | Pass | Pass | Pass
SST-06 | Security | 1. Execute Export in Native (Raw) Format | File contents are encrypted and stored. | File contents are encrypted and stored. | Pass | Pass | Pass
SST-07 |  | 1. Download Log File from the Log Screen | All user actions are recorded in the log file. | All user actions are recorded in the log file. | Pass | Pass | Pass
SST-08 |  | 1. Enter ID/PW on the Login Screen | The program runs only when the correct ID/PW is entered. | The program runs only when the correct ID/PW is entered. | Pass | Pass | Pass
SST-09 |  | 1. Execute Wired/Wireless Network Settings | KIOSK mode is enabled, restricting network setting functionalities. | KIOSK mode is enabled, restricting network setting functionalities. | Pass | Pass | Pass