Software V&VT report

Product: Rudens

Model: Faster

Document No.: VVT-01

|  | PREPARE | REVIEW | APPROVE |
| --- | --- | --- | --- |
| Signature |  |  |  |
| Date | 2022.01.15 | 2022.01.17 | 2022.01.17 |

Revision History

| 제/개정 번호 | 제/개정 일자 | 주요 제/개정 내용 |
| --- | --- | --- |
| 1.0 | 2023.07.28 | 최초 개정 |

# 1. 적용 범위 및 목적 (Scope and Purpose)

## 1.1 적용범위 (Purpose)

본 문서는 소프트웨어 verification 및 validation 결과에 대한 설명을 제공한다.

## 1.2 적용범위 (Scope)

VVT 문서는 MCU의 소프트웨어 verification 및 validation에만 적용된다. 하드웨어에 대한 내용은 포함되어 있지 않는다.

## 1.3 소프트웨어 식별 (Software Identification)

| 품목군 | 생체현상 측정기기 |  |  |
| --- | --- | --- | --- |
| 품목명 | 의료영상저장전송분석장치 |  |  |
| 모델명 | Faster |  |  |
| 소프트웨어 명칭 | Rudens | 소프트웨어 버전 | 1.00.00 |
| 소프트웨어 안전등급 | Class B |  |  |

## 1.4 관련 규격 (Reference)

| Reference | Title of standard |
| --- | --- |
| 1) ISO 14971 [2007] | Medical devices –Application of risk management to medical devices. |
| 2) IEC 60601-1 [2005]+AM1[2012] | Medical electrical equipment – General requirements for basic safety and essential performance |
| 3) IEC 60601-1-2 [2014] | Medical electrical equipment. General requirements for basic safety and essential performance. Collateral Standard. Electromagnetic disturbances. Requirements and tests. |
| 4) ISO 7010:2011 | Graphical symbols -- Safety colours and safety signs -- Registered safety signs |
| 5) IEC 60601-1-6 [2010] +AMD1 [2013] | Medical Electrical Equipment - Part 1-6: General Requirements for Basic Safety and Essential Performance - Collateral Standard: Usability |
| 6) IEC 62366 Ed. 1.0 [2007] +AMD1[2014] | Medical devices - Application of usability engineering to medical devices |
| 7) 의료기기 법, 의료기기법 시행규칙 | 의료기기법 및 같은 법 시행령에서 위임된 사항과 그 시행에 필요한 사항을 규정 |
| 8) 의료기기제조 및 품질관리기준 | 제조 또는 수입함에 있어 준수하여야 하는 세부사항과 품질관리심사기관 및 품질책임자 교육실시기관에 대한 지정절차 및 관리방법 등에 관하여 필요한 사항을 정함 |

# 2. 테스트 계획 (Test plan)

# 2.1. 일반 시험내용

| 테스트 아이디 (Test ID) | DT1~6 (통합/시스템 테스트), 정적 테스트는 Tool 에서 생성하는 Tag 로 대체 |
| --- | --- |
| 소프트웨어 관련 정보 | Rudens 1.00.00 |
| 소프트웨어 테스트 유형 | 정적/통합/시스템 테스트 |

# 2.2 준비 사항

| 하드웨어 준비 | - 장비 준비 및 전원 On - 카테터 준비 및 Rotary Junction에 연결 |
| --- | --- |
| 소프트웨어 준비 | SW 실행 |

# 3. 테스트 결과 (Test result)

# 3.1 통합/시스템 테스트 결과 (절차 및 적합/부적합 기준, 관계된 요구사항)

| Test case | Item | Method | Criteria | SRS |
| --- | --- | --- | --- | --- |
| DT1 | Initialize |  |  |  |
| DT1-1 | Initialize Acquisition Board | 1. OCT Scanner Mode | 프로그램과 Acquisition Device가 연결되었는지 확인 | SRS-17 |
|  |  | 1-1. Start Program |  |  |
|  |  | 1-2. Click Initialize Button |  |  |
|  |  | 2. Administrator Mode |  |  |
| DT1-2 | Initialize Rotary Junction | 2-1. Start Program | 프로그램과 Motor가 연결되었는지 확인 | SRS-02 SRS-17 |
| DT1-3 | Control Linear Stage | 2-2. Click Admin Button | 연결된 Linear Stage 의 위치가 초기화되었는지 확인 | SRS-14 |
| DT1-5 | Laser Power On | 2-3. Click Initialize Button | 연결된 Laser 전원이 켜졌는지 확인 | SRS-30 SRS-31 |
| DT1-4 | Indicate Current State | 1. Start Program | 프로그램의 좌측 상단에 상태 메시지가 출력되는지 확인 | SRS-01 |
|  |  | 2. Click Initialize Button |  |  |
| DT2 | OCT Image |  |  |  |
| DT2-1 | Invert Image | 1. Start Program | 클릭시 OCT 영상이 반전되는지 확인 | SRS-03 SRS-07 SRS-18 SRS-20 |
|  |  | 2. Click Admin Button |  |  |
|  |  | 3. Click Initialize Button |  |  |
|  |  | 4. Select Black / White Options |  |  |
| DT2-2 | Apply Color Map | 1. Start Program | 활성화시 색상을 포함한 Image를 출력 비활성화시 흑백 Image 를 출력 | SRS-03 SRS-07 SRS-18 SRS-20 |
|  |  | 2. Click Admin Button |  |  |
|  |  | 3. Click Hot Color Checkbox |  |  |
| DT2-3 | Apply Brightness, Contrast | 1. OCT Scanner Mode | 설정값에 따라 OCT 영상의 밝기와 채도가 변화하는지 확인 | SRS-04 SRS-05 SRS-28 SRS-29 |
|  |  | 1-1. Start Program |  |  |
|  |  | 1-2. Click Initialize Button |  |  |
|  |  | 1-3. Slide Brightness, Contrast Bar |  |  |
|  |  | 2. Administrator Mode |  |  |
|  |  | 2-1. Start Program |  |  |
|  |  | 2-2. Click Admin Button |  |  |
|  |  | 2-3. Click Initialize Button |  |  |
|  |  | 2-4. Slide Brightness, Contrast Bar |  |  |
| DT2-4 | Display Scope | 1. Start Program | OCT 신호 및 FFT 처리된 신호가 화면에 출력되는지 확인 | SRS-03 SRS-07 SRS-18 |
|  |  | 2. Click Admin Button |  |  |
|  |  | 3. Click Initialize Button |  |  |
|  |  | 4. Select Scope Radio Button |  |  |
| DT2-5 | Display Circle Image | 1. Start Program | 원형 OCT 영상이 화면에 출력되는지 확인 | SRS-03 SRS-07 SRS-18 |
|  |  | 2. Click Admin Button |  |  |
|  |  | 3. Click Initialize Button |  |  |
|  |  | 4. Select Circle Radio Button |  |  |
| DT2-6 | Display Rectangle Image | 1. Start Program | 사각형 OCT 영상이 화면에 출력되는지 확인 | SRS-03 SRS-07 SRS-18 SRS-20 |
|  |  | 2. Click Admin Button |  |  |
|  |  | 3. Click Initialize Button |  |  |
|  |  | 4. Select Rectangle Radio Button |  |  |
| DT3 | Save Data |  |  |  |
| DT3-1 | Input Patient Name | 1. Start Program 2. Click Initialize Button 3. Type Name Into PATIENT NAME 4. Click Scan Button | Patient의 이름이 정상적으로 입력되는지 확인 이름을 입력하지 않을 경우, 촬영이 시작되지 않는지 확인 | SRS-06 |
| DT3-2 | Save OCT Data |  | Linear Stage Pullback 이 끝난 후, 메모리에 적재된 OCT 신호가 파일로 저장되었는지 확인 | SRS-16 SRS-25 |
| DT4 | Device Control |  |  |  |
| DT4-1 | Rotary Junction Setting | 1. OCT Scanner Mode | Rotary Junction 설정값이 해당 메뉴에 표시된다 | SRS-12 SRS-13 |
|  |  | 1-1. Start Program | 변경한 설정값은 저장되어 이후에도 프로그램에 반영된다. |  |
|  |  | 1-2. Click Settings Button |  |  |
|  |  | 2. Administrator Mode |  |  |
|  |  | 2-1. Start Program |  |  |
|  |  | 2-2. Click Admin Button |  |  |
|  |  | 2-3. Click Rotary Junction Open Button |  |  |
| DT4-2 | Move Linear Stage | 1. OCT Scanner Mode | Linear Stage가 초기 위치나 UI 에 입력된 위치, 속도로 이동하는지 확인 | SRS-14 |
|  |  | 1-1. Start Program |  |  |
|  |  | 1-2. Click Initialize Button |  |  |
|  |  | 1-3. Click Settings Button |  |  |
|  |  | 1-4. Click Idle / Pull / Move Buttons |  |  |
|  |  | 2. Administrator Mode |  |  |
|  |  | 2-1. Start Program |  |  |
|  |  | 2-2. Click Admin Button |  |  |
|  |  | 2-3. Click Initialize Button |  |  |
|  |  | 2-4. Click Rotary Junction Open Button |  |  |
|  |  | 2-5. Click Idle / Pull / Move Buttons |  |  |
| DT4-3 | Move Interferometer Stage | 1. OCT Scanner Mode | Interferometer Stage가 UI 입력에 따라 앞뒤로 이동하는지 확인 | SRS-15 |
|  |  | 1-1. Start Program |  |  |
|  |  | 1-2. Click Initialize Button |  |  |
|  |  | 1-3. Click Settings Button |  |  |
|  |  | 1-4. Click Forward / Backward Buttons |  |  |
|  |  | 2. Administrator Mode |  |  |
|  |  | 2-1. Start Program |  |  |
|  |  | 2-2. Click Admin Button |  |  |
|  |  | 2-3. Click Initialize Button |  |  |
|  |  | 2-4. Click Rotary Junction Open Button |  |  |
|  |  | 2-5. Click Forward / Backward Buttons |  |  |
| DT4-4 | Motor Control | 1. OCT Scanner Mode | Motor 가동시 설정한 속도로 Motor가 회전해야 한다 | SRS-09 SRS-16 |
|  |  | 1-1. Start Program |  |  |
|  |  | 1-2. Click Initialize Button |  |  |
|  |  | 1-3. Click Motor On Button |  |  |
|  |  | 2. Administrator Mode |  |  |
|  |  | 2-1. Start Program |  |  |
|  |  | 2-2. Click Admin Button |  |  |
|  |  | 2-3. Click Initialize Button |  |  |
|  |  | 2-4. Click Rotary Junction Open Button |  |  |
|  |  | 2-5. Click Perform Run Button |  |  |
| DT4-5 | Shutter Control | 1. Start Program | 설정값에 따라 Shutter가 동작(Open/Close)하는지 확인 | SRS-23 |
|  |  | 2. Click Admin Button |  |  |
|  |  | 3. Click Close Shutter Checkbox |  |  |
| DT5 | Simulation |  |  |  |
| DT5-1 | Simulate Scanned OCT Image | 1. OCT Scanner Mode | OCT 영상을 재생하거나 일시정지 후 이전, 다음 프레임으로 이동한다 | SRS-08 SRS-11 SRS-24 |
|  |  | 1-1. Start Program |  |  |
|  |  | 1-2. Click Initialize Button |  |  |
|  |  | 1-3. Type Name Into PATIENT NAME |  |  |
|  |  | 1-4. Click Scan Button |  |  |
|  |  | 1-5. Click Play / Pause / Previous / Next Button |  |  |
|  |  | 2. Administrator Mode |  |  |
|  |  | 2-1. Start Program |  |  |
|  |  | 2-2. Click Admin Button |  |  |
|  |  | 2-3. Select File |  |  |
|  |  | 2-4. Click Load Button |  |  |
|  |  | 2-5. Click Play Button |  |  |
| DT5-2 | Display Cut-view Image | 1. Start Program | Linear Stage Pullback 이 끝난 후, OCT 영상 하단에 촬영된 OCT영상으로부터 단면 영상이 생성되는지 확인 | SRS-24 |
|  |  | 2. Click Initialize Button |  |  |
|  |  | 3. Type Name Into PATIENT NAME |  |  |
|  |  | 4. Click Scan Button |  |  |
| DT5-3 | Save Video File | 1. Start Program | 화면에 출력되는 OCT 영상이 동영상 파일로 저장되는지 확인 | SRS-26 |
|  |  | 2. Click Admin Button |  |  |
|  |  | 3. Select File |  |  |
|  |  | 4. Click Load, Play Button |  |  |
|  |  | 5. Click Save AVI Button |  |  |
| DT6 | Save Calibration Data |  |  |  |
| DT6-1 | Save Calibration Data | 1. Start Program | Linear Stage 가 입력한 위치로 이동하면서 OCT 기기로부터 수신한 데이터가 파일로 저장되는지 확인 | SRS-27 |
|  |  | 2. Click Admin Button |  |  |
|  |  | 3. Click Initialize Button |  |  |
|  |  | 4. Type from / step / count into edit text |  |  |
|  |  | 5. Click Save Button |  |  |

# 3.2 정적 테스트 결과

| Checker |  | Passed? | Remark |
| --- | --- | --- | --- |
| TAG | Summary | (Y/N) |  |
| C1255 | PCH data for plugin has incorrect length. | Y | The Code Analysis tool reports this warning when there is an internal error in the tool, not in the code being analyzed. |
| C1256 | PCH must be both written and read. | Y | The Code Analysis tool reports this warning when there is an internal error in the tool, not in the code being analyzed. |
| C6001 | using uninitialized memory <variable> | Y | This warning is reported when an uninitialized local variable is used before it is assigned a value. This could lead to unpredictable results. You should always initialize variables before use. |
| C6011 | dereferencing NULL pointer <name> | Y | This warning indicates that your code dereferences a potentially null pointer. If the pointer value is invalid, the result is undefined. To resolve the issue, validate the pointer before use. |
| C6014 | Leaking memory. | Y | This warning indicates that the specified pointer points to allocated memory or some other allocated resource that has not been freed. The analyzer checks for this condition only when the _Analysis_mode_(_Analysis_local_leak_checks_) SAL annotation is specified. By default, this annotation is specified for Windows kernel mode (driver) code. |
| C6029 | possible buffer overrun in call to <function>: use of unchecked value | Y | This warning indicates that a function that takes a buffer and a size is being passed a unchecked size. The data read-in from some external source has not been verified to see whether it is smaller than the buffer size. An attacker might intentionally specify a much larger than expected value for the size, which will lead to a buffer overrun. |
| C6031 | return value ignored: called-function could return unexpected value | Y | This warning indicates the caller doesn't check a function's return value for failure. Depending on which function is being called, this defect can lead to seemingly random program misbehavior. That includes crashes and data corruptions in error conditions or low-resource situations. |
| C6053 | call to <function> may not zero-terminate string <variable> | Y | This warning indicates that the specified function has been called in such a way that the resulting string might not be zero-terminated. This defect might cause an exploitable buffer overrun or crash. This warning is also generated if an annotated function expects a null terminated string is passed a string that is not null terminated. |
| C6054 | string <variable> may not be zero-terminated | Y | This warning indicates that a function that requires a zero-terminated string was passed a non-zero terminated string. A function that expects a zero-terminated string could look for the zero beyond the end of the buffer. This defect might cause an exploitable buffer overrun error or crash. The program should make sure the string passed in ends with a zero. |
| C6059 | Incorrect length parameter in call to <function>. Pass the number of remaining characters, not the buffer size of <variable> | Y | This warning indicates that a call to a string concatenation function is probably passing an incorrect value for the number of characters to concatenate. This defect might cause an exploitable buffer overrun or crash. A common cause of this defect is passing the buffer size, instead of the remaining number of characters in the buffer, to the string manipulation function. |
| C6063 | missing string argument to <function> corresponding to conversion specifier <number> | Y | This warning indicates that not enough arguments are being provided to match a format string; at least one of the missing arguments is a string. This defect can cause crashes and buffer overflows (if the called function is of the sprintf family), as well as potentially incorrect output. |
| C6064 | missing integer argument to <function> corresponding to conversion specifier <number> | Y | This warning indicates that not enough arguments are being provided to match a format string and one of the missing arguments is an integer. This defect can cause incorrect output. |
| C6066 | non-pointer passed as parameter <number> when pointer is required in call to <function> | Y | This warning indicates that the format string specifies that a pointer is required, for example, a %n or %p specification for printf or a %d for scanf, but a non-pointer is being passed. This defect is likely to cause a crash or corruption of some form. |
| C6067 | parameter <number> in call to <function> must be the address of the string | Y | This warning indicates a mismatch between the format specifier and the function parameter. Even though the warning suggests using the address of the string, you must check the type of parameter a function expects before correcting the problem. For example, a %s specification for printf requires a string argument, but a %s specification in scanf requires an address of the string. |
| C6101 | Returning uninitialized memory | Y | A successful path through the function does not set the named _Out_ parameter. This message is generated based on SAL annotations that indicate that the function in question always succeeds. |
| C6102 | Using <variable> from failed function call at line <location>. | Y | This warning is reported instead of C6001 when a variable is not set because it was marked as an _Out_ parameter on a prior function call that failed. |
| C6103 | Returning <variable> from failed function call at line <location>. | Y | A successful path through the function is returning a variable that was used as an _Out_ parameter to an internal function call that failed. |
| C6200 | index <name> is out of valid index range <min> to <max> for non-stack buffer <variable> | Y | This warning indicates that an integer offset into the specified array exceeds the maximum bounds of that array. This defect might cause random behavior or crashes. |
| C6201 | buffer overrun for <variable>, which is possibly stack allocated: index <name> is out of valid index range <min> to <max> | Y | This warning indicates that an integer offset into the specified stack array exceeds the maximum bounds of that array. This defect might cause random behavior or crashes. |
|  |  |  | One common cause of this defect is using an array's size as an index into the array. Because C/C++ array indexing is zero-based, the maximum legal index into an array is one less than the number of array elements. |
| C6294 | Ill-defined for-loop: initial condition does not satisfy test. Loop body not executed | Y | This warning indicates that a for-loop cannot be executed because the terminating condition is true. This warning suggests that the programmer's intent is not correctly captured. |
| C6386 | buffer overrun: accessing <buffer name>, the writable size is <size1> bytes, but <size2> bytes may be written: Lines: x, y | Y | This warning indicates that the writable extent of the specified buffer might be smaller than the index used to write to it. This can cause buffer overrun. |
| C26451 | Using operator 'operator' on a size-a byte value and then casting the result to a size-b byte value. Cast the value to the wider type before calling operator 'operator' to avoid overflow | Y | Code analysis detects when an integral value gets shifted left, multiplied, added, or subtracted, and the result gets cast to a wider integral type. If the operation overflows the narrower integral type, then data is lost. You can prevent this loss by casting the value to a wider type before the arithmetic operation. |
| C26454 | Arithmetic overflow: '%operator%' operation produces a negative unsigned result at compile time | Y | This warning indicates that the subtraction operation produces a negative result which was evaluated in an unsigned context. This can result in unintended overflows. |
| C26495 | MEMBER_UNINIT | Y | Variable '%variable%' is uninitialized. Always initialize a member variable |  |
| C26812 | Prefer 'enum class' over 'enum' (Enum.3) | Y | The enum type type-name is unscoped. Prefer 'enum class' over 'enum' (Enum.3) |  |