## Traceability
SRS ID | SDS ID | Module | Function | Functional requirements | SUT ID | Module | Function | Description | SIT ID | Item | Verification | Expected Result | Actual Result | SST ID | Item | Test Procedure | Expected Result | Actual Result | SRM ID
SRS-001 | SDS-001 | UI Display | Page UI | Page display | SUT-01 | UI Display | void Navigate(object parameter) | Page display | SIT-01 | Page display | Check if the Patient List screen is displayed after loading | Patient List screen is displayed. | Patient List screen is displayed. | SST-01 | UI Display | 1. Access Home (Patient List) Screen 2. Execute Import Function | The main screen and pop-ups are displayed. | The main screen and pop-ups are displayed. | SRM-001
SRS-002 | SDS-002 |  | Dialog UI | Dialog display | SUT-02 |  | DialogResults OpenDialog(object dialog, Dictionary<string, object> parameter, double parentWidth, double parentHeight, double left, double top) | Dialog display |  | Dialog display | Check if the Import Dialog (Popup) appears when clicking the Import button | Import Popup is displayed. | Import Popup is displayed.
SRS-003 | SDS-003 | DB Processing | Database Insert | Data input into Database | SUT-03 | DB Processing | int InsertData(string commandText, Dictionary<string, Object> commandParameters = null); | Data input into Database | SIT-02 | Data input into Database | Check if a new patient is added via New Patient | A new patient is added. | A new patient is added. | SST-02 | DB Processing | 1. Add new patient information in the Add Patient screen 2. Modify patient information in the Edit Patient screen 3. Delete patient information in the Edit Patient screen 4. Check patient information on the Patient List screen | Patient information can be added, modified, deleted, and viewed. | Patient information can be added, modified, deleted, and viewed. | SRM-002
SRS-004 | SDS-004 |  | Database Select | Data retrieval from Database | SUT-04 |  | IList<T> GetDatas<T>(string commandText, Dictionary<string, Object> commandParameters = null) where T : class; | Data retrieval from Database |  | Data retrieval from Database | Check if the patient list is displayed on the Patient List screen | The patient list is displayed on the Patient List screen. | The patient list is displayed on the Patient List screen.
SRS-005 | SDS-005 |  | Database Update | Data modification in Database | SUT-05 |  | int UpdateData(string commandText, Dictionary<string, Object> commandParameters = null); | Data modification in Database |  | Data modification in Database | Check if patient information is modified in Edit Patient | Patient information is modified. | Patient information is modified.
SRS-006 | SDS-006 |  | Database Delete | Data deletion in Database | SUT-06 |  | int DeleteData(string commandText, Dictionary<string, Object> commandParameters = null); | Data deletion in Database |  | Data deletion in Database | Check if the patient is removed from the list after deletion | Patient information is deleted. | Patient information is deleted.
SRS-007 | SDS-007 | Rotary Junction Control | Control BLDC Motor | BLDC Motor rotation control | SUT-07 | Rotary Junction Control | bool CMotorController::PerformRun(int &nVelocity) | BLDC Motor rotation control | SIT-03 | BLDC Motor rotation control | Check if the Catheter connector rotates when connected to RJ | A red light is displayed at the tip of the Catheter. | A red light is displayed at the tip of the Catheter. | SST-03 | Catheter Connection | 1. Configure recording settings on the Recording Settings screen and proceed to Next 2. Connect RJ to Catheter 3. Press the Live button on the Live View screen 4. Execute Auto Calibration on the Calibration screen | OCT images are displayed, and OCT image Calibration is executed. | OCT images are displayed, and OCT image Calibration is executed. | SRM-003 SRM-004 SRM-005
SRS-008 | SDS-008 |  | Control Pullback Motor | Pullback Motor control to adjust Linear Stage | SUT-08 |  | bool CRJController::Move(eStepMotorIndex idxMotor, int posStep, bool delay, char sensor) | Pullback Motor control to adjust Linear Stage |  | Pullback Motor control to adjust Linear Stage | Check if the Catheter connector moves when connected to RJ | A red light is displayed at the tip of the Catheter. | A red light is displayed at the tip of the Catheter.
SRS-009 | SDS-009 |  | Check Catheter Connection Status | Acquisition of Catheter connection status | SUT-09 |  | UINT COCTSystem::threadValidateCatheter(LPVOID param) | Acquisition of Catheter connection status |  | Acquisition of Catheter connection status | Check if VLD turns on at the tip of the Catheter | A red light is displayed at the tip of the Catheter. | A red light is displayed at the tip of the Catheter.
SRS-010 | SDS-010 | Laser Module Control | Control Visible Light | Visible Light power control | SUT-10 | Laser Module Control | void CLaserModule::SetVLD(unsigned short vld) | Visible Light power control | SIT-04 | Visible Light power control | Check VLD at the tip of the Catheter | A red light is displayed at the tip of the Catheter. | A red light is displayed at the tip of the Catheter.
SRS-011 | SDS-011 |  | Control Laser | Laser power control | SUT-11 |  | int CLaserController::LaserOnOff(bool on) | Laser power control |  | Laser power control | Check if real-time OCT images are displayed on the LiveView screen | OCT image is displayed. | OCT image is displayed.
SRS-012 | SDS-012 |  | Control Delay Line Motor | Delay Line Motor control to adjust Linear Stage | SUT-12 |  | bool CLaserModule::Move(eStepMotorIndex idxMotor, int posStep, bool delay, char sensor) | Delay Line Motor control to adjust Linear Stage |  | Delay Line Motor control to adjust Linear Stage | Check if Calibration works on the LiveView screen | Calibration is activated. | Calibration is activated.
SRS-013 | SDS-013 |  | Control Polarization Motor | Polarization Motor control for angle adjustment | SUT-13 |  | bool CLaserModule::Move(eStepMotorIndex idxMotor, int posStep, bool delay, char sensor) | Polarization Motor control for angle adjustment |  | Polarization Motor control for angle adjustment | Check if Calibration works on the LiveView screen | Calibration is activated. | Calibration is activated.
SRS-014 | SDS-014 | Acquisition Control | Acquire OCT Signal | Acquisition of OCT Signal from Acquisition Board | SUT-14 | Acquisition Control | UINT IAcquisitionDevice::threadAcquire(LPVOID param) | Acquisition of OCT Signal from Acquisition Board | SIT-05 | Acquisition of OCT Signal from Acquisition Board | Check if real-time OCT images are displayed on the LiveView screen | OCT image is displayed. | OCT image is displayed. | SST-04 | Recording & Review | 1. Recording 1-1. Start Recording in the Live View screen 1-2. Execute Ready and Start in the Recording screen 1-3. Select Accept and Continue on the Image Quality Check screen 2. Review (2D) 2-1. Activate OCT mode in the Review (2D) screen 2-2. Activate OCT + Angio mode in the Review (2D) screen 2-3. Activate Measure-Auto in the Review (2D) screen 2-4. Activate Measure-Manual and perform measurement in the Review (2D) screen 2-5. Operate Presentation in the Review (2D) screen 2-6. Activate Lumen Profile in the Review (2D) screen 2-7. Activate L-View in the Review (2D) screen 2-8. Adjust Frame/Degree Indicator in the Review (2D) screen 3. Review (3D) 3-1. Change 3D Option in the Review (3D) screen 3-2. Adjust Frame/Degree Indicator in the Review (3D) screen (Note: The displayed content may vary depending on Pre/Post/Follow-up cases.) | 1. Recording Screen - OCT and Angio images are recorded. 2. Review (2D) Screen - Cross-Section, Angio, Lumen Profile, and L-View are displayed. - Lumen Contour, Side Branch, Stent, and Calcium information are displayed. - Automatically measured values such as Area, Diameter, %AS, %DS, Expansion, and Apposition are displayed. - Manually measured values such as Area, Diameter, Length, and Annotations are displayed. - Cross-Section and L-View are configured according to Indicator and Presentation. 3. Review (3D) Screen - 3D Objects are displayed in Cut-View and Fly-Through formats. - Objects are selectively displayed based on option selection. - Cut-View and Fly-Through are configured according to the Indicator. | 1. Recording Screen - OCT and Angio images are recorded. 2. Review (2D) Screen - Cross-Section, Angio, Lumen Profile, and L-View are displayed. - Lumen Contour, Side Branch, Stent, and Calcium information are displayed. - Automatically measured values such as Area, Diameter, %AS, %DS, Expansion, and Apposition are displayed. - Manually measured values such as Area, Diameter, Length, and Annotations are displayed. - Cross-Section and L-View are configured according to Indicator and Presentation. 3. Review (3D) Screen - 3D Objects are displayed in Cut-View and Fly-Through formats. - Objects are selectively displayed based on option selection. - Cut-View and Fly-Through are configured according to the Indicator. | SRM-006 SRM-007 SRM-008 SRM-009 SRM-010 SRM-011 SRM-012 SRM-013
SRS-015 | SDS-015 | Signal Processing | Convert OCT Signal to Image | Conversion of OCT Signal to Image format | SUT-15 | Signal Processing | void COCTImaging::Process(char* fringes) | Conversion of OCT Signal to Image format |  | Conversion of OCT Signal to Image format | Check if real-time OCT images are displayed on the LiveView screen | OCT image is displayed. | OCT image is displayed.
SRS-016 | SDS-016 |  | Load Raw Image | Loading Raw Image file | SUT-16 |  | RayError COCTSystem::OpenImage(char* strFilePath, double imageResolution, double zOffset) | Loading Raw Image file |  | Loading Raw Image file | Check if OCT images are displayed in Review | OCT image is displayed. | OCT image is displayed.
SRS-017 | SDS-017 |  | Save Raw Image | Saving Raw Image file | SUT-17 |  | UINT COCTSystem::threadSaveRaw(LPVOID param) | Saving Raw Image file |  | Saving Raw Image file | Open a saved Case to check the OCT image | OCT image is displayed. | OCT image is displayed.
SRS-018 | SDS-018 |  | Make L-View Image | Construction of L-View Image from Cross-Section Image | SUT-18 |  | UINT CImagingSession::threadUpdateCutView(LPVOID param) | Construction of L-View Image from Cross-Section Image |  | Construction of L-View Image from Cross-Section Image | Check L-View in Review | L-View is displayed. | L-View is displayed.
SRS-019 | SDS-019 |  | Change Degree | Degree adjustment of OCT Image to rotate L-View Image | SUT-19 |  | RayError COCTSystem::SetDegree(double value) | Degree adjustment of OCT Image to rotate L-View Image |  | Degree adjustment of OCT Image to rotate L-View Image | Check if L-View rotates when adjusting the Degree Indicator in Review | L-View rotates while being displayed. | L-View rotates while being displayed.
SRS-020 | SDS-020 |  | Change Frame | Frame position adjustment to modify Cross-Section Image | SUT-20 |  | void* COCTSystem::GetImageData(int nFrame) | Frame position adjustment to modify Cross-Section Image |  | Frame position adjustment to modify Cross-Section Image | Check if Cross-Section Image changes when adjusting the Frame Indicator in Review | Cross-Section image is updated. | Cross-Section image is updated.
SRS-021 | SDS-021 |  | Adjust Brightness | OCT Image Brightness adjustment | SUT-21 |  | RayError COCTSystem::SetBrightness(double value) | OCT Image Brightness adjustment |  | OCT Image Brightness adjustment | Check brightness adjustment in Review | Brightness of the image is adjusted. | Brightness of the image is adjusted.
SRS-022 | SDS-022 |  | Adjust Contrast | OCT Image Contrast adjustment | SUT-22 |  | RayError COCTSystem::SetContrast(double value) | OCT Image Contrast adjustment |  | OCT Image Contrast adjustment | Check contrast adjustment in Review | Contrast of the image is adjusted. | Contrast of the image is adjusted.
SRS-023 | SDS-023 | ML Processing | Detect Lumen Contour | Extraction of Lumen Contour information from OCT Image | SUT-23 | ML Processing | cv::Mat CRayYolo::FindLumen(cv::Mat image) | Extraction of Lumen Contour information from OCT Image | SIT-06 | Extraction of Lumen Contour information from OCT Image | Check Lumen Contour information after Recording | Lumen Contour is displayed in the Cross-Section. | Lumen Contour is displayed in the Cross-Section.
SRS-024 | SDS-024 |  | Detect Side Branch | Extraction of Side Branch information from OCT Image | SUT-24 |  | cv::Mat CRayYolo::FindSidebranch() | Extraction of Side Branch information from OCT Image |  | Extraction of Side Branch information from OCT Image | Check Side Branch information after Recording | Side Branch location is displayed in the Lumen Profile. | Side Branch location is displayed in the Lumen Profile.
SRS-025 | SDS-025 |  | Detect Stent | Extraction of Stent information from OCT Image | SUT-25 |  | std::vector<cv::Rect2f> CRayYolo::FindStent(cv::Mat image) | Extraction of Stent information from OCT Image |  | Extraction of Stent information from OCT Image | Check Stent information after Recording | Stent location is displayed in the Cross-Section. | Stent location is displayed in the Cross-Section.
SRS-026 | SDS-026 |  | Detect Guidewire | Extraction of Guidewire information from OCT Image | SUT-26 |  | std::vector<cv::Rect2f> CRayYolo::FindGuidewire() | Extraction of Guidewire information from OCT Image |  | Extraction of Guidewire information from OCT Image | Check Guidewire information after Recording (3D screen) | Guidewire is displayed in the Cut-View of the 3D image. | Guidewire is displayed in the Cut-View of the 3D image.
SRS-027 | SDS-027 |  | Detect Calcium | Extraction of Calcium information from OCT Image | SUT-27 |  | void* RayGetCalciumAngles(int nFrame) | Extraction of Calcium information from OCT Image |  | Extraction of Calcium information from OCT Image | Check Calcium information after Recording | Calcium location is displayed in the Cross-Section. | Calcium location is displayed in the Cross-Section.
SRS-028 | SDS-028 | 3D Rendering | 3D Cut-View Rendering | 3D Cut-View image rendering | SUT-28 | 3D Rendering | static extern int ODSOCT_Render(); | 3D Cut-View image rendering | SIT-07 | 3D Cut-View image rendering | Check Cut-View images in the 3D screen | 3D Object is displayed in the Cut-View area. | 3D Object is displayed in the Cut-View area.
SRS-029 | SDS-029 |  | 3D Fly-Through Rendering | 3D Fly-Through image rendering | SUT-29 |  | static extern int ODSOCT_Render(); | 3D Fly-Through image rendering |  | 3D Fly-Through image rendering | Check Fly-Through images in the 3D screen | 3D Object is displayed in the Fly-Through area. | 3D Object is displayed in the Fly-Through area.
SRS-030 | SDS-030 |  | Rotate 3D Object | 3D Object rotation | SUT-30 |  | static extern int ODSOCT_RotateAngle(float angle); | 3D Object rotation |  | 3D Object rotation | Check 3D Object rotation when adjusting the Degree Indicator in the 3D screen | 3D Object rotates while being displayed. | 3D Object rotates while being displayed.
SRS-031 | SDS-031 |  | Translate 3D Camera | Camera Object position adjustment | SUT-31 |  | static extern int ODSOCT_MoveCameraPosition(int direction, bool inverse); | Camera Object position adjustment |  | Camera Object position adjustment | Check 3D Object position adjustment when manipulating the Frame Indicator in the 3D screen | Position of the 3D Object is updated. | Position of the 3D Object is updated.
SRS-032 | SDS-032 |  | Change 3D Object Rendering Option | Modification of Tissue/Lumen/Guidewire/Side Branch rendering method | SUT-32 |  | static extern int ODSOCT_SetViewData(Ray3DObject obj, Ray3DObjectMode mode); | Modification of Tissue/Lumen/Guidewire/Side Branch rendering method |  | Modification of Tissue/Lumen/Guidewire/Side Branch rendering method | Check rendering changes when modifying 3D screen options | 3D Object display is updated based on the selected Option. | 3D Object display is updated based on the selected Option.
SRS-033 | SDS-033 | Measurement | Measure Contour | Measurement of Area, Diameter, %AS, %DS | SUT-33 | Measurement | void DrawLumenContour(LumenContour lumenContour, LumenStent lumenStent, double appositionThreshold, bool isEditOn) | Measurement of Area, Diameter, %AS, %DS | SIT-08 | Measurement of Area, Diameter, %AS, %DS | Check Lumen information (Area, Diameter, %AS, %DS) in Review | Lumen information (Area, Diameter, %AS, %DS) is displayed. | Lumen information (Area, Diameter, %AS, %DS) is displayed.
SRS-034 | SDS-034 |  | Measure Stent | Measurement of Expansion, Apposition | SUT-34 |  | void DrawLumenContour(LumenContour lumenContour, LumenStent lumenStent, double appositionThreshold, bool isEditOn) | Measurement of Expansion, Apposition |  | Measurement of Expansion, Apposition | Check Stent information (Expansion, Apposition) in Review | Stent information (Expansion, Apposition) is displayed. | Stent information (Expansion, Apposition) is displayed.
SRS-035 | SDS-035 |  | Measure Area | Measurement of Area and Diameter of a specific region | SUT-35 |  | void DrawCurve(AreaGeometry areaGeometry, bool isCurrentEditOn = true) | Measurement of Area and Diameter of a specific region |  | Measurement of Area and Diameter of a specific region | Check Area and Diameter by drawing a region in Review | An area is added to the image, and Area, Length information is displayed. | An area is added to the image, and Area, Length information is displayed.
SRS-036 | SDS-036 |  | Measure Length | Measurement of Length between two points | SUT-36 |  | void DrawLine(Point firstPoint, Point secondPoint, int group, bool isEnd) | Measurement of Length between two points |  | Measurement of Length between two points | Check Length by drawing a line in Review | A line is added to the image, and Length information is displayed. | A line is added to the image, and Length information is displayed.
SRS-037 | SDS-037 |  | Make an Annotation | Annotation creation | SUT-37 |  | void DrawText(Point point, int group, string text) | Annotation creation |  | Annotation creation | Check annotation display in Review | Annotations are added to the image. | Annotations are added to the image.
SRS-038 | SDS-038 | Angio Client | Control Angio Server | Angio Server process control | SUT-38 | Angio Client | ConnectionStatus ConnectToServer() | Angio Server process control | SIT-09 | Angio Server process control | Check if the Angio icon is activated in the bottom left | Angio icon is activated in the bottom left corner. | Angio icon is activated in the bottom left corner.
SRS-039 | SDS-039 |  | Get Angio Image | Receiving Angio Image from Angio Server | SUT-39 |  | void ThreadFuncLiveAngioImage() | Receiving Angio Image from Angio Server |  | Receiving Angio Image from Angio Server | Check Angio images in the Recording screen | Real-time Angio video is displayed. | Real-time Angio video is displayed.
SRS-040 | SDS-040 |  | Load Angio Raw Image | Loading Angio Raw Image file | SUT-40 |  | void ThreadReadAngioFrames() | Loading Angio Raw Image file |  | Loading Angio Raw Image file | Check if Angio images are displayed in Review | Angio video is displayed. | Angio video is displayed.
SRS-041 | SDS-041 |  | Save Angio Raw Image | Saving Angio Raw Image file | SUT-41 |  | void ThreadFuncSaveAngioFrames(PatientCase patientCase) | Saving Angio Raw Image file |  | Saving Angio Raw Image file | Open a saved Case to check the Angio image | Angio video is displayed. | Angio video is displayed.
SRS-042 | SDS-042 | Angio Server | Acquire Angio Image | Acquiring Angio Image from Angio equipment | SUT-42 | Angio Server | void TCPSocket::LiveFrameThread(FrameGrabber& fg) | Acquiring Angio Image from Angio equipment |  | Acquiring Angio Image from Angio equipment | Check Angio images in the Recording screen | Real-time Angio video is displayed. | Real-time Angio video is displayed.
SRS-043 | SDS-043 |  | Send Angio Image | Transmitting acquired Angio Image to Angio Client | SUT-43 |  | void TCPSocket::LiveFrame(FrameGrabber& fg) | Transmitting acquired Angio Image to Angio Client |  | Transmitting acquired Angio Image to Angio Client | Check Angio images in the Recording screen | Real-time Angio video is displayed. | Real-time Angio video is displayed.
SRS-044 | SDS-044 | Image Processing | Make Lumen Profile | Generating Lumen Profile Image using Lumen Contour information | SUT-44 | Image Processing | bool DrawLumenProfileImage() | Generating Lumen Profile Image using Lumen Contour information | SIT-10 | Generating Lumen Profile Image using Lumen Contour information | Check Lumen Profile in Review | Lumen Profile is displayed. | Lumen Profile is displayed. | SST-05 | Export | 1. Execute Export Function 2. Select Export Type and the Patient Case to Extract 3. Choose Options according to Each Export Type and Execute Export | Images are generated according to the selected Export Type and Option settings. | Images are generated according to the selected Export Type and Option settings. | SRM-014
SRS-045 | SDS-045 |  | Make Export Image File | Generating Review screen as an Export Image | SUT-45 |  | static async Task<List<Mat>> MakeImageForExport(PatientCase patientCase, List<Mat> imgCrossSections, Mat imgLongitude, List<int>? exportIndices, FileExport fileExport, Action<double> progressCallback, double progress, Action<string> progressTextCallback) | Generating Review screen as an Export Image |  | Generating Review screen as an Export Image | Check image results in Export | The image is displayed according to the selected Export Option. | The image is displayed according to the selected Export Option.
SRS-046 | SDS-046 |  | Convert OCT Image to Standard Format | Converting Export Image to Standard format (MP4, JPEG, Bitmap, TIFF) | SUT-46 |  | async Task FileSaveStandard() | Converting Export Image to Standard format (MP4, JPEG, Bitmap, TIFF) |  | Converting Export Image to Standard format (MP4, JPEG, Bitmap, TIFF) | Check saved format in Export results | Exported files are saved in the corresponding format. | Exported files are saved in the corresponding format.
SRS-047 | SDS-047 | DICOM Processing | Convert OCT Image to DICOM Format | Converting Export Image to DICOM format | SUT-47 | DICOM Processing | async Task FileSaveDicom() | Converting Export Image to DICOM format |  | Converting Export Image to DICOM format | Check stored DICOM type in Export results | Exported files are saved in DICOM format. | Exported files are saved in DICOM format.
SRS-048 | SDS-048 | Security | Data Security | Encrypting important information during data extraction | SUT-48 | Security | static async Task Encryptor(string filePath, string contents, Action<double> progressCallback, double progressSize, Action<string> progressTextCallback) | Encrypting important information during data extraction | SIT-11 | Encrypting important information during data extraction | Check if file contents are encrypted in Export results | Exported file contents are encrypted. | Exported file contents are encrypted. | SST-06 | Security | 1. Execute Export in Native (Raw) Format | File contents are encrypted and stored. | File contents are encrypted and stored. | SRM-015
SRS-049 | SDS-049 |  | Logging | Logging functionality for access, query, creation, modification, and deletion | SUT-49 |  | void Debug(object message) | Logging functionality for access, query, creation, modification, and deletion | SIT-12 | Logging functionality for access, query, creation, modification, and deletion | Check Log file content by retrieving Log files from the Setting menu | Log files record access, query, creation, modification, and deletion history. | Log files record access, query, creation, modification, and deletion history. | SST-07 |  | 1. Download Log File from the Log Screen | All user actions are recorded in the log file. | All user actions are recorded in the log file. | SRM-016
SRS-050 | SDS-050 |  | System Access | Only authorized users can access the system by entering ID/PW | SUT-50 |  | Functions Outside of Proprietary Software (Windows Functions) | Only authorized users can access the system by entering ID/PW | SIT-13 | Only authorized users can access the system by entering ID/PW | Check if system access requires user ID/PW entry | User must enter the correct ID/PW to access. | User must enter the correct ID/PW to access. | SST-08 |  | 1. Enter ID/PW on the Login Screen | The program runs only when the correct ID/PW is entered. | The program runs only when the correct ID/PW is entered. | SRM-017
SRS-051 | SDS-051 |  | Data Communication Access Control | Only USB data communication is allowed, with wired and wireless communication restricted | SUT-51 |  | Functions Outside of Proprietary Software (Windows Functions) | Only USB data communication is allowed, with wired and wireless communication restricted | SIT-14 | Only USB data communication is allowed, with wired and wireless communication restricted | Check if the device has a LAN port and supports Wi-Fi functionality | The device has no LAN port, and Wi-Fi functionality is restricted. | The device has no LAN port, and Wi-Fi functionality is restricted. | SST-09 |  | 1. Execute Wired/Wireless Network Settings | KIOSK mode is enabled, restricting network setting functionalities. | KIOSK mode is enabled, restricting network setting functionalities. | SRM-018