## SRS
SRS ID | Module | Function | Functional requirements
| UI Display | Page UI | Page display
|  | Dialog UI | Dialog display
| DB Processing | Database Insert | Data input into Database
|  | Database Select | Data retrieval from Database
|  | Database Update | Data modification in Database
|  | Database Delete | Data deletion in Database
| Rotary Junction Control | Engage catheter to rotary junction interface | When the user connects the catheter to the rotary junction, the system engages the rotary junction interface to the catheter.
|  | Verify catheter | The system verifies that the catheter is from an approved manufacturer and that the image is valid after connection.
|  | Control BLDC motor for rotary junction | After catheter connection, the system rotates the BLDC motor so that the user can check the intravascular image.
|  | Pullback function | The system rotates the rotary junction and catheter and adjusts their position so that the inside of the vessel can be imaged.
|  | Disconnect catheter from rotary junction | After image acquisition is completed, the system disconnects the catheter from the rotary junction interface.
|  | Emergency stop function | In case of abnormal operation or malfunction, the system can stop the rotary junction.
|  | Rotary junction status indication | The system displays the current status of the rotary junction to the user.
|  | Fiber Cleaning Mode | The system moves the rotary junction interface forward so that the fiber can be cleaned.
| Laser Module Control | VLD on/off | After catheter connection, the system turns on the VLD red light so that the user can confirm the connection status.
|  | VOA on/off | After catheter connection, the system turns on the VOA light so that the user can confirm the image.
|  | Control Delay Line Motor | After catheter connection, the system drives the Delay Line Motor to perform image calibration.
|  | Control Polarization Motor | After catheter connection, the system drives the Polarization Motor to perform image calibration.
| Acquisition Control | Acquire OCT signal | After catheter connection, the system commands the signal processing board to acquire the laser signal as OCT image data.
|  | Stop acquiring OCT signal | When the catheter is disconnected or image acquisition must be stopped, the system commands the signal processing board to stop data acquisition.
| Signal Processing | Set Laser Module Setting | Sets laser acquisition parameters such as sampling interval, A-scan/B-scan size, and trigger options.
|  | Load System Calibration File | Loads background, k-linearization map, and dispersion data from a specified path and applies them to processing.
|  | Get OCT Signal from Laser Module | Acquires OCT fringe signals from the laser module at predefined intervals using current hardware settings.
|  | Remove Noise Signal from OCT Signal | Removes DC background and reduces speckle/other noise to improve OCT signal quality before imaging.
|  | Convert OCT Signal to Image | Converts processed OCT signals to cross-sectional images via FFT, interpolation, and intensity mapping.
|  | Delay Line Calibration | Adjusts the optical delay line so that image depth position is calibrated and remains stable.
|  | Polarization Calibration | Adjusts system polarization to maximize signal intensity and reduce polarization-related artifacts.
|  | Load Raw Image | Loads previously saved raw OCT image or signal data from storage into memory.
|  | Save Raw Image | Saves current raw OCT image or signal data to storage in the selected file format.
|  | Make L-View Image | Builds an L-View (longitudinal) image by stacking sequential cross-sectional OCT frames.
|  | Change Degree | Rotates the displayed OCT or L-View image by a specified angle for better viewing.
|  | Change Frame | Changes the currently displayed frame within a pullback sequence for navigation along the vessel.
|  | Adjust Brightness | Adjusts overall image brightness (intensity level) to improve visibility.
|  | Adjust Contrast | Adjusts image contrast to enhance differences between bright and dark regions.
| ML Processing | Load ML Model | Loads and initializes a pre-trained ML model from local storage for automatic detection tasks.
|  | Detect Lumen Contour | Automatically extracts the lumen boundary from the OCT image as points or a mask.
|  | Detect Side Branch | Detects side branches in the OCT image and extracts their locations and shapes.
|  | Detect Stent | Detects stent struts/structure in the OCT image and records their positions for analysis.
|  | Detect Guidewire | Detects the guidewire and extracts its path so it can be displayed or masked.
|  | Detect Calcium | Detects calcified regions in the OCT image and extracts their location and extent.
|  | Draw Overlay | Draws object overlays (lumen, stent, side branch, guidewire, calcium, etc.) on the OCT image.
| 3D Rendering | 3D Cut-View Rendering | The system shall generate a 3D cut-view model from OCT images.
|  | 3D Cut-View Half-Section Display | The system shall display half-section 3D cut surfaces in the 3D cut-view.
|  | 3D Cut-View Zoom Control | The system shall allow zoom-in and zoom-out on the 3D cut-view model.
|  | 3D Fly-Through Rendering | The system shall generate a 3D fly-through rendering from OCT images.
|  | 3D Fly-Through Direction Control | The system shall allow changing the viewing direction (distal-to-proximal / proximal-to-distal).
|  | Rotate 3D Object | The system shall allow users to rotate the 3D object by dragging a rotation control.
|  | Camera Rotation Sync | The system shall update all linked views (cut-view, fly-through, and L-view) based on camera rotation.
|  | Translate 3D Camera | The system shall move the 3D camera to the position corresponding to the selected frame of the L-View.
|  | Camera Translation Sync | The system shall update all linked views (cut-view, fly-through, and 2D OCT view) based on camera translation.
|  | Current Frame Indicator Toggle | The system shall allow enabling or disabling the camera position indicator.
|  | 3D Rendering Toggle | The system shall allow enabling or disabling Tissue, Lumen, Stent, Guidewire, Side Branch rendering in 3D mode.
| Measurement | Measure Contour | Measurement of Area, Diameter, %AS, %DS
|  | Measure Stent | Measurement of Expansion, Apposition
|  | Measure Area | Measurement of Area and Diameter of a specific region
|  | Measure Length | Measurement of Length between two points
|  | Measure Angle | Measurement of Angle between three points
|  | Make an Annotation | Annotation creation
| Angio Client | Control Angio Server | Angio Server process control
|  | Get Angio Image | Receiving Angio Image from Angio Server
|  | Load Angio Raw Image | Loading Angio Raw Image file
|  | Save Angio Raw Image | Saving Angio Raw Image file
| Angio Server | Acquire Angio Image | Acquiring Angio Image from Angio equipment
|  | Send Angio Image | Transmitting acquired Angio Image to Angio Client
| Image Processing | Lumen Profile Generation | The system shall generate a lumen profile image using the detected lumen contour information.
|  | Review Image Export | The system shall generate export images based on the current review screen.
|  | Standard Format Conversion | The system shall convert export images into standard formats, including MP4, JPEG, Bitmap, and TIFF.
|  | Manual Calibration | The system shall allow users to adjust the global Z-offset applied to all OCT frames.
|  | Review Calibration | The system shall adjust the frame alignment based on the detected sheath position in each OCT frame.
|  | Image Compensation | The system shall apply noise reduction to OCT images.
| DICOM Processing | Convert OCT Image to DICOM Format | Converting Export Image to DICOM format
|  | Retrieve Patient Information | Retrieve Patient Information from the DICOM Server (PACS/MWL)
|  | DICOM Transfer | Transmit DICOM Data to PACS
| Security | Data Extraction Security | The system shall encrypt important information during extraction.
|  | Database Encryption | The system shall encrypt all important information stored in the internal database.
|  | System Logging | The system shall record logs for user access, configuration changes, data retrieval, data creation, data modification, and data deletion events, including the timestamp for each logged event.
|  | Audit Log Integrity | The system shall protect audit logs from unauthorized modification or deletion.
|  | System Access Control | The system shall allow access only to authenticated users using valid ID and password credentials.
|  | Data Communication Access Control | The system shall restrict data communication to USB-only interfaces and disable wired and wireless network communication.
|  | Account Role Management | The system shall implement role-based access control to separate permissions among different user roles.
|  | Security Settings Access Control | The system shall restrict access to security-related configuration settings to authorized administrator roles only.
|  | Session Timeout | The system shall automatically log out inactive users after a predefined timeout period.
|  | Password Policy Enforcement | The system shall enforce a password policy that includes minimum length and complexity requirements.
|  | Password Expiration Notification | The system shall notify users when their password is expired or requires updating.
|  | Failed Login Attempt Delay | The system shall introduce a delay after a predefined number of consecutive failed login attempts.
|  | Malware Protection | The system shall provide protection against malicious software by restricting unauthorized program execution.
|  | Data Backup and Recovery | The system shall provide a secure data backup and recovery mechanism to prevent loss of patient information.
|  | Storage Capacity Monitoring | The system shall notify the user with a warning message when the available storage capacity falls below a predefined threshold.